## Slide 1
EPAC ISTAC
Detectors and
Diagnostics
C. D. Armstrong, D. Symes, J. Green 14/05/2024
Notes: The abstract pattern can be removed or repositioned if required. Be careful to ‘Send to Back’ so that it does not obscure any important information.

## Slide 2
Overview
Overview
Project Scope
Diagnostics for LWFA (EA1)
  Standard experimental requirements
  In-area laser and optical diagnostics
  Electron spectrometer and beamline
  X-ray characterisation
Diagnostics for Solid targets (EA2)
  Standard experimental requirements
  Ion spectral/spatial diagnostics
  X-ray characterisation
  Neutron detectors
Notes: Outline the scope/challenge of the project and then talk about what the key radiation types we expect from each area in EPAC to give context to the detector designs I’m going to show.
I’ll run through the concepts we’ve come up with for each detector and then run through where our technical risks are and how they’ve changed since last year.
Then to wrap up I’ll show off some of the prototyping results in a few areas and talk about what it means going forward for each design.

## Slide 3: Project Scope

## Slide 4
What is required?
X-rays:
10 keV to 100 MeV detection
Source-limited spatial resolution at 100 keV
Efficient detection at high energies
Protons / Ions:
1 to 100 MeV/u detection
Active replacement for RCF (spectro-spatial imaging)
High dynamic range spectrometer
Neutrons:
Epithermal to Fast detection
100 ps temporal resolution
Discretion from x-rays
Electrons:
<100 MeV (EA2) to 10 GeV (EA1)
Charge, emissivity, spectral characterisation all required to high quality
Wide angular/spatial information for EA2
J. Won et al. Nucl. Fusion 63 (2023) 066031
M. Schollmeier, et al. Rev. Sci. Instrum. 1 April 2014; 85 (4): 043305
C. D. Armstrong, et al. Front. Phys. (11) (2023): 28-30.
A. R. Maier, et al. PRX 10.3 (2020): 031039.

## Slide 5
What is being delivered?
A baseline suite of detectors and diagnostics able to characterise standard experimental configurations and industrial applications.
The critical diagnostics to perform a wide variety of experiments.
Distinction here is a detector is sensitive to the radiation and the diagnostic is a total package.
* to some degree – resolution limits or energies requirements might not necessarily cover specific applications
This requires sufficient understanding and development to establish.
X-rays:
High spatial resolution <100 keV
Wide FOV (>150mm) imaging across energy domains
High efficiency 100 keV imaging
~keV resolution spectroscopy up to 600 keV
Spectral fitting up to 100s MeV
Protons / Ions:
<100 MeV Thomson spectroscopy
3-Layer spectro-spatial monitor as “active-RCF”
Neutrons:
Time gated imaging with fast CH scintillator
High temporal resolution nTOF
Electrons:
Compact <100 MeV spectroscopy
% level 0.5-10 GeV spectroscopy
Beam pointing/distribution monitors in line
Non-invasive charge measurements
Longitudinal characterisation via CTR

## Slide 6
Electron
Probe
Exit Mode
Electron
Transverse
Longitudinal
Whole beam
Wavefront
Scintillator screens
Permanent double dipole
Quad emittance
Non-invasive
charge measurement
Variable electromagnet
Dielectric Streaker
Transmission laser spectroscopy
Imaging
Optics
Detector
Underpinning Methodology
Scintillator
Spectroscopy
Exit Mode
Magnets
EA1

## Slide 7
Protons / Ions
EA2
Optical
Imaging
Optics
Detector
Underpinning Methodology
Scintillator
Spectroscopy
Protons / Ions
Scintillator screens
Three colour splitting system
Electrons
Optical
Transverse probe
Transmission/reflection
Magnets
Active Thomson parabola
Isolated chamber
Compact design
Compact active e-spec
Laser harmonics

## Slide 8
Imaging
Optics
Detector
Underpinning Methodology
Betatron
Particle
Hard X-ray
Neutron
Scintillator
High Rep-Activation
X-ray
THz
High Flux CZT
Electron Cherenkov
2D Active Thomson
Spectroscopy
Multi-Modal
Bayer Ross-Pair
Linear Spectrometer
Interferometer CTR
Splitting Spectrometer
N-/P-
Shaped
Scintillators
MeV+ Imager
Fast Decay
Sub-micron imager
Large area panel
Detectors
Directly coupled
Low noise system

## Slide 9: Experimental Requirements for EA1
Predominantly gas targets with long focal length parabolas. Narrow emission angle limits opportunity for simultaneous measurements. Requires characterisation of ultra-relativistic electron beam, large energy spread of x-rays, and high resolution spatial profiles.

## Slide 10
EA1 Examples Experimental Layout
Experiment description:
Gas jet / cell interaction with long focus
Sustained 10 Hz operation with variable gas density
Investigating the coupling of laser-energy into the wakefield
Expected variation from wavefront and gas-density/structure
Collimated emission area with predominately electron and x-ray emission
Diagnostic Areas:
Laser input conditions
Laser-plasma coupling characterisation
Plasma density
Emitted electron beam characteristics
keV to MeV x-ray emission

## Slide 11
Exit mode characterisation
Direct
Indirect
T. Sokollik, et al., AIP Conf. Proc. 1299, 233 (2010).
Provide the users with on-shot:
Transmitted energy fraction (NF integral)
Spectral modulation (Shamrock)
Pointing/Hosing (FF centroid)
Wavefront (Zernike modes)
A.J. Ross et al. https://arxiv.org/pdf/2310.05097
Develop towards:
SW-/MW- Infrared Spectroscopy
Indirect plasma mirror operation for compact exit mode configuraitons

## Slide 12
Target density and profile
Provide the users with on-shot:
Interferometry of time-integrated plasma channels across mm- and um- scales
30 fs probe at 1st and 2nd harmonic
Analysis platform matured to “gui-interface” of probe density reconstruction
Develop towards:
Sub-10fs probes to open up wakefield structure directly
SW-/MW- IR probe to reduce the plasma density characterisation below 1017
Chirped co-linear probe increases sensitivity to lower average plasma density
Schwab et al PRAB 23, 032802 (2020)
Detailed design ongoing
Increasing complexity
Transverse probe

## Slide 13
Electron beam emission - Spectra
Provide the users with:
Calibrated and variable dipole for electron characterisation.
Target electron energies returned to consistent position and nominally colinear propagation below axis.
Double variable strength dipole for flexible operation
Develop towards:
Calibrated and variable dipole for electron characterisation.
Target electron energies returned to consistent position and nominally colinear propagation below axis.

## Slide 14
Electron beam emission – Spatial
Provide the users with on-shot:
Spatial profile of electron beam emission at fixed distances from target
Invasive measurement with scintillator
Multiple units to provide degree of emissivity measurements
Develop towards:
Non-invasive measurements for passive characterisation / stability monitor
High resolution scanning system (requires source stability)
Thinfilm optical CTR characterisation to minimise in-beam material

## Slide 15
Electron beam emission – Longitudinal
Provide the users with THz CTR emission:
Interferometry across 1-30 um with dual view
Temporal FOV out to ~400fs in single frame
Form-factor pulse reconstruction of complex profiles
Broadband THz spectroscopy
Develop towards:
Ultra-high sensitivity single pixel detectors
Compressed sensing algorithms to enhance reconstruction of low signal
See poster tomorrow!

## Slide 16
X-ray Characterisation – HE Spatial (10-300 keV)
Provide the users with on-shot:
Highly efficiency detection of <100 keV x-ray.
Blurring PSF equivalent to 20um at the scintillator plane
Masking techniques to provide on-shot / in-image spectral reconstruction
Large Area: 300x300 mm at ~400 um resolution
High resolution: 16 x 14 mm at ~50 um resolution
J. Wood, ‘Betatron radiation from laser wakefield accelerators and its applications’, Sep. 2016, doi: 10.25560/58282

## Slide 17
X-ray Characterisation – HR Spatial (10-300 keV)
Provide the users with on-shot:
Dual view at both ~50 and ~4 um plane resolution limits
Macro-view acts as a viewfinder for operation a mitigates any shot-shot instability.
Micro-view under few-x geometric magnification hits sub-micron resolving power of x-rays up to 100 keV.
Develop towards:
High repetition scanning of micro-view arm
Variable scintillator options
Larger FOV, lower noise.
Simulated radiograph
Macro View
Micro View

## Slide 18
X-ray Characterisation – Spectral (10-600 keV)
Provide the users with on-shot:
76k px x-ray spectrometer with ~keV resolution to 600 keV
Target operation at 15% population to provide onshot characterisation of the full spectra.
Indirect/scattering methods to reduce flux and retain spectral information
Develop towards:
Spatial spectral techniques using multi-pinhole arrangements (multiple shots)
Diffractive measurements using combination of displacement and energy
High energy calibrations for Compton backscatter
Veale, M. C., Seller, P., Wilson, M., & Liotti, E. (2018). Synchrotron Radiation News, 31(6), 28–32. https://doi.org/10.1080/08940886.2018.1528431
See poster tomorrow!

## Slide 19
X-ray Characterisation – Spectral (0.3-300 MeV)
LWFA Bremsstrahlung
Inverse Compton Scatter
K. Behm et al. Rev. Sci. Instrum. 89, 113303 (2018)
Broadband spectral characterisation:
For higher energies we need significantly more attenuation
Past users have used ~300mm of CsI (4g/cc)
Applied a similar reconstruction process to before iteratively fitting the expected shape with the response matrix

## Slide 20
X-ray Characterisation – Spectral (0.3-300 MeV)
LWFA Bremsstrahlung
Inverse Compton Scatter
Broadband spectral characterisation:
For higher energies we need significantly more attenuation
Past users have used ~300mm of CsI (4g/cc)
Applied a similar reconstruction process to before iteratively fitting the expected shape with the response matrix
EPAC version is ~250 mm of LYSO (7g/cc) split across 3 camera systems with independent gain and lens systems to maximise dynamic range
Camera 2
Camera 3
Camera 1

## Slide 21: Experimental Requirements for EA2
Predominantly solid targets with short focal length parabolas. Requires characterisation of emitted particle and photon radiation, emission is simultaneous/co-linear so isolation methods must be considered.

## Slide 22
EA2 Examples Experimental Layout
Experiment description:
Solid target interaction using short focal length parabola at high intensity
Structured tape targets spooled at 1 Hz operation
Investigating the coupling of laser-energy into the accelerated electrons and secondary radiation
Expected variation from focus, pre-plasma formation, laser/target parameters
Large emission area with mix of radiation types
Diagnostic Areas:
Laser input conditions
Laser-plasma coupling characterisation
Target evolution
Emitted proton/ion radiation
X-ray radiation
Neutrons from secondary converter (ion/x-ray source)

## Slide 23
Emitted proton/ion distribution - spatial
Provide the users with on-shot:
Develop towards:
11.3MeV
5.8 MeV
9.1 MeV
Provide the users:
Few energy-bin system for inspecting spatial proton/ion emission
RCF-like at 10 Hz
“Probies” style mask to provide further attenuation regions

## Slide 24
Source
Emitted proton/ion distribution - spectral
Provide the users with on-shot:
2x dedicated Thomson parabola angles
Characterised ”TAP” style active Thomson with high brightness phosphor screen
Proton/ion distinction with variable e-field
Develop towards:
2D array pinhole development for ~degree resolution proton characterisation
Source

## Slide 25
Emitted neutron distribution - spatial
Provide the users:
2-4x 10 ns gated images of fast scintillator system
Spatially resolving isolated neutron energies up to 100 MeV
Gamma-discretion through time rather than PSD
Develop towards:
Armstrong et al. Submitted to Review of Scientific Instruments, (2024)

## Slide 26
Emitted neutron distribution - spectral
Provide the users:
~200 ps temporal resolution of arriving signal
Neutron and proton sensitivity
3x channels to provide multiple angles of sampling
Temporal response
High sensitivity response
Hyperbolic shape redirects signal to PMT
High reflectivity coating to maintain intensity
<ns response, volume limited
Blackened surfaces only permit direct light to PMT
Kills “whispering gallery”
~100ps response material limited
Scintillator
PMT
Develop towards:
<100 ps temporal resolution
In-vacuum mounting for p-TOF response
Current version
Pushing to faster responses

## Slide 27
X-ray characterisation - spatial
Provide the users with on-shot:
100-400 um resolution over 150 x 150 mm2 active FOV
~10% absorbed energy at 10 MeV (for 5 mm thick system)
1 mm thick system equivalent sensitivity to commercial FOP with approximately 3x better resolution
5 mm thick system approximately an order of magnitude higher sensitivity (E>1 MeV) with equivalent resolution

## Slide 28
X-rays:
High spatial resolution <100 keV
Wide FOV (>150mm) imaging across energy domains
High efficiency 100 keV imaging
~keV resolution spectroscopy up to 600 keV
Spectral fitting up to 100s MeV
Protons / Ions:
<100 MeV Thomson spectroscopy
3-Layer spectro-spatial monitor as “active-RCF”
Neutrons:
Time gated imaging with fast CH scintillator
High temporal resolution nTOF
Electrons:
Compact <100 MeV spectroscopy
% level 0.5-10 GeV spectroscopy
Beam pointing/distribution monitors in line
Electron beamline optics
Non-invasive charge measurements
Longitudinal characterisation via CTR
Summary
Providing a comprehensive suite of detectors and diagnostics to cover base needs from the user/industrial community
Laser input conditions, plasma/target evolution, and emitted radiation diagnostics.
Highlights:
  Sub-micron system for 100 keV betatron
  <200 um system for 1-10 MeV x-rays
  Longitudinal and emissivity diagnostics for electron optimisation

## Slide 29: Extra Slides

## Slide 30
Betatron Emission
Transition lines (k-..)
Bremsstrahlung
ICS
2ndary Nuclear Emissions
Positron Annihilation
DLA Induced X-rays
HHG
T. Wang, Phys. Rev. Applied 13, 054024 (2020)
T. Popmintchev, et al. Science 336.6086 (2012)
R. J., Murphy, et al. Astrophys. Journal Supp. Series 161.2 (2005)
M.M Günther, et al. Nat Commun 13, 170 (2022).
H-E Tsai, et al. Physics of Plasmas 22.2 (2015).
F Mackenroth, et al. Physical Review E 99.3 (2019): 033205.
P. Neumayer, et al. Physics of plasmas 17.10 (2010).
S. Cipiccia, et al. Journal of Applied Physics 111.6 (2012).
R. D. Edwards, et al. Applied Physics Letters 80.12 (2002)
F. Albert, and A. G. R. Thomas. Plasma Physics and Controlled Fusion 58.10 (2016)
X-ray sources
Notes: The abstract pattern can be removed or repositioned if required. Be careful to ‘Send to Back’ so that it does not obscure any important information.

## Slide 31
Compton Scattering
Pair Production
Transmissive
Transmission/Absorption Spectroscopy
Activation
Single Hit Spectroscopy
Compton Recoil
Cascade Detection
Kojima et al. Rev. Sci. Instrum. 87, 043502 (2016)
Reflective
Bragg Crystal Refraction
Ross-Pair Differential Spectroscopy
Talbot/Fresnel Grating Spectroscopy
C. Svetina, et al. Opt. Lett. 44, 574-577 (2019)
J Seely. RSI. 93, 093529 (2022)
K. W. Hill et al. RSI. 87, 11E344 (2016)
Maddox, B. R., et al. RSI 82.2 (2011).
Chen, C. D., et al. RSI 79.10 (2008).
Rusby, D. R., et al. RSI 89.7 (2018).
Behm, K. T., et al. RSI 89.11 (2018).
Rakowski et al. Scientific REPORTS | 7: 16603 (2017)
R. J. Protheroe, MNRAS 221.4 (1986): 769-788
M. I. K. Santala, et al. APL 78.1 (2001)
Photo-electric
X-ray Characterisation – Spectral Techniques
Notes: The abstract pattern can be removed or repositioned if required. Be careful to ‘Send to Back’ so that it does not obscure any important information.

## Slide 32
X-ray Characterisation – Spectral (10-300 keV)
Provide the users with on-shot:
Spectral characterisation between ~10 keV and 500 keV critical energy
<% resolution (for betatron spectral shape) across energy range
”At-rep” analysis routine for incident single temperature
~Hz level analysis routine for complex spectral shapes
Develop towards:
”At-rep” analysis routine for incident complex spectra
“Anomaly” detection for “unknown” spectral features
ML-reconstruction
Optimised scintillator patterns
Broadband spectral characterisation:
Triple axis linear absorption spectrometer
High efficiency measure of incident radiation with minor angle () between each carriage.
Either:
  Angular variation from fixed pattern offset
  Spectral region change when
Synthetic data
180 keV
180 keV
200 keV

## Slide 33
Laser characterisation diagnostics
Laser monitor via leaked through final turning mirror
Highly reflective for CW and alignment
AR coating at 800 for full power
NF/FF for stabilisation/ pilot beam
Wavefront monitor and NF/FF for main beam
Wedge to minimise ghosts/secondary surface reflections
Provide the users with on-shot:
Laser energy (NF integral)
Beam footprint (NF uniformity)
Pointing (FF centroid)
Wavefront (Zernike modes)
From upstairs:
Pulse duration (SPIDER)
  Pick off reference to whole beam characterisation
SOD, TOD terms
Spectrometer

## Slide 34
Electron beam emission – Emissivity
Provide the users with ability to:
Vary quadrupole position and monitor impact on the spatial profile of the beam
Develop towards:
Active methods for focus and emissivity scanning of quadrupoles as routine.
10x permanent quad set for long-distance focus
Floor dictated by plasma mirror
Permanent high-field quadrupole to capture/control electron beam
Monitor (2)
Quad (1)

## Slide 35
Laser characterisation diagnostics
Laser monitor via leaked through final turning mirror
Highly reflective for CW and alignment
AR coating at 800 for full power
Wavefront monitor and NF/FF for main beam
Wedge to minimise ghosts/secondary surface reflections
Provide the users with on-shot:
Laser energy (NF integral)
Beam footprint (NF uniformity)
Pointing (FF centroid)
Wavefront (Zernike modes)
Spectrometer
NF/FF for stabilisation/ pilot beam

## Slide 36
Escaping electron distribution
Active compact e-spec design:
Field-tested design with configurations up to 50 MeV
Small-throw, double-magnet, reduced field versions tested with community.
Provide the users:
3-4x chassis with variable internals
Calibrated and known configurations to work up to 100 MeV
Develop towards:
Spectro-spatial developments with coded-aperture/complex pinhole

## Slide 37
X-ray Characterisation – Spectral (10-600 keV)
Provide the users with on-shot:
76k px x-ray spectrometer with ~keV resolution to 600 keV
Target operation at 15% population to provide onshot characterisation of the full spectra.
Indirect/scattering methods to reduce flux and retain spectral information
Develop towards:
Spatial spectral techniques using multi-pinhole arrangements (multiple shots)
Diffractive measurements using combination of displacement and energy
High energy calibrations for Compton backscatter
Veale, M. C., Seller, P., Wilson, M., & Liotti, E. (2018). Synchrotron Radiation News, 31(6), 28–32. https://doi.org/10.1080/08940886.2018.1528431
See poster tomorrow!

## Slide 38
X-ray Characterisation – Spectral (0.3-10 MeV)
Provide the users with on-shot:
Spectral characterisation between ~100 keV and 10 MeV Boltzmann temperatures
<5% resolution across energy range
”At-rep” analysis routine for incident single temperature
~Hz level analysis routine for complex spectral shapes
Develop towards:
”At-rep” analysis routine for incident complex spectra
“Anomaly” detection for “unknown” spectral features
ML-reconstruction
Optimised scintillator patterns
Broadband spectral characterisation:
  Triple axis linear absorption spectrometer
  High efficiency measure of incident radiation with minor angle () between each carriage.
  Either:
    Angular variation from fixed pattern offset
    Spectral region change when
Synthetic data
180 keV
180 keV
200 keV

## Slide 39
Concept
Design
Prototyping
Procurement
Characterisation
Commissioning
Development
Spectroscopy
Imaging
Sub-micron,
<100 keV x-ray
MeV Imager,
MeV x-ray
Multi-modal,
>100 keV x-ray, N
Bayer Mask
<100 keV x-ray
LAS,
MeV x-ray
nTOF
< 100 MeV Neutron
THz
0.1-50 Thz
Status at 2021 ISTAC

## Slide 40
Concept
Design
Prototyping
Procurement
Characterisation
Commissioning
Development
Spectroscopy
Imaging
Sub-micron,
<100 keV x-ray
MeV Imager,
MeV x-ray
Multi-modal,
>100 keV x-ray, N
Bayer Mask
<100 keV x-ray
LAS,
MeV x-ray
nTOF
< 100 MeV Neutron
THz
0.1-50 Thz
Large Area Panel,
High sensitivity
Status at 2022 ISTAC

## Slide 41
Concept
Design
Prototyping
Procurement
Characterisation
Commissioning
Development
Spectroscopy
Imaging
Sub-micron,
<100 keV x-ray
MeV Imager,
MeV x-ray
Multi-modal,
>100 keV x-ray, N
Bayer Mask
<100 keV x-ray
LAS, (3 versions)
MeV x-ray
nTOF
< 100 MeV Neutron
THz
0.1-50 THz
Large Area Panel,
High sensitivity
Status at 2022 ISTAC
Direct Coupled,
Efficient/Resolution, x
Hexitec
10-600 keV x-ray