## Slide 1
Design of an achromatic multi-pass relay-imaging telescope for the EPAC high-energy Ti:Sa amplifier
Tiago de Faria Pinto, Rob Heathcote, Mark Harman, Paul Mason, Cristina Hernandez-Gomez
Central Laser Facility, Science and Technology Facilities Council, Rutherford Appleton Laboratory,
Harwell Science and Innovation Campus, Didcot OX11 0QX, United Kingdom
Abstract
Thermal lens simulations
EPAC multi-pass relay imaging
Custom doublet design
Future work
Thermal lens compensation
Challenges:
Focusing back reflections create a plasma in air which compromises subsequent pulses
Tilting the lenses to deflect these reflections severely degrades performance
Low n2 requirements restrict glass choices
Solution:
Double meniscus doublets
  No focusing 1st order reflections
  Still has focusing reflections from vacuum windows, but those can be tilted without a performance drop
  Second-order ghosts enclosed in vacuum
Peak-to-valley wavefront error after 4 passes below λ
  Well within adaptive optic correction
4-pass chromatic aberration below λ/10
Bint below 0.12 for all telescopes combined
Input and output doublets are the same to reduce costs
VSF input doublet of each pass moves equal distance to compensate thermal lens changes
Large changes in thermal lens
  Caused by changes to pump energy or repetition rate
  Lenses move to preset positions for coarse correction
Small changes in thermal lens
  Caused by changes to pump-seed timing and day-to-day variation
  Lenses finely optimised and wavefront optimised with an adaptive optic
4-pass cooled Ti:Sa amplifier
1:1 relay telescope between each pass, 3 m focal length
Thermal lens is compensated at each pass to ensure proper collimation and wavefront at the crystal
Vacuum spatial filters to mitigate back-propagating reflections
Characterisation and management of focusing back-reflections
Thermal lens created on the Ti:Sa crystal has been simulated for various pumping energies and cooling conditions:
  Thermal changes to refractive index
  Bulk thermal expansion
  Stress-induced birefringence
Thermal lens simulations incorporated in the OpticStudio models used to design the imaging system
System requirements
Bint below 0.5 for the entire multi-pass section
Chromatic wavefront error below λ/10
Wavefront error with adaptive optic compensation range
Needs to accommodate large changes in thermal lens
Careful management of focusing reflections
  Eliminate 1st order focusing reflections
  2nd order focusing reflections away from optics
Ti:Sa amplifier head
Moveable doublets
Vacuum spatial filters
N-LAK22
S-FPL53
Without simulated AO correction
With simulated AO correction
Lenses move 90 mm between cold system and maximum thermal lens (full pumping, no extraction)
  Doublet position automated
  by the control system
  Specifications met over the
  full doublet position range
  Ghost reflections characterised
  over the full movement range
N-LAK22
S-FPL53
June 2024: Optics installation
July 2024: CW alignment
September 2024: Low-energy pulsed alignment
October 2024: High-energy - 1 Hz
November 2024: High-energy - 10 Hz
Laser
Notes: Print-ready version
The colours in this template have been adjusted for printing purposes. Use this template if the primary use for your poster will be in print.