## Slide 1
This work was supported in part by the Swiss National Science Foundation
E-mail: alexander.karpushov@epfl.ch WWW: spc.epfl.ch
Diagnostics for characterization of neutral beams parameters at TCV
A.N. Karpushov, Y. Andrebe, B.P. Duval, R. Jacquier, A. Listopad, D. Mykytchuk, H. Reimerdes, M. Toussaint, and TCV teama
Ecole Polytechnique Federale de Lausanne (EPFL), Swiss Plasma Center (SPC), CH-1015 Lausanne, Switzerland
aSee author list of B.P. Duval et al 2024 Nucl. Fusion 64 112023, DOI:10.1088/1741-4326/ad8361
21st International Conference on Ion Sources, Oxford, UK, 8th–12th of September 2025 ID:#10
NBIs on TCV
Multi-channel spectroscopy
calorimeter
3.2 m
aiming CEMs
optical port
cryo-panels
magnet
ion dump
neutralizer
RF plasma
source
ion optics
(grids)
TCV[1]: RO/a 0.88/0.25m BT≤1.54T
NBI-1 ≤1.3MW, ≤28keV (D/H) &
NBI-2 ≤1.1MW, ≤52keV (D/H)
RF plasma source, triode IOS with slit beamlets, geometrical focusing
“standard” set of diagnostics: grids U&I, RF power, magnet I, TCs on calorimeter and ion dump, vacuum, …
DNBI 80kW, 50keV (H)
cold cathode arc-plasma source, tetrode IOS with circular beamlets
NBI-2 layout (similar NBI-1) [2]
Beam duct protection: NBI-2/1 max.280/150MW/m2 − 17×22cm2 ports
Tokamak protection: shine-through on central column and beam dump
NBI profiles (F(r3,v3)) essential for:
  Optimization of beam duct – min. losses, overheat, beam blocking
  Modelling of NBI deposition in plasma (fast ion birth rate)
  Interpretation of CXRS measurement – DNBI profiles in plasma
NBI-1/2 ducts
Design optimization:
  Minimal power density on walls
  Minimal surface temperature
  Surface material – heat remove & low sputtering, release of gas/impurities
CuCrZr beam duct insert[3]
250 
156×206mm2
Measurements: 19 thermocouples & optical (350-1100 nm) detector
Tungsten target
Measurement of 3-D NB power density distribution
  4mm actively cooled tungsten tile at 45° to beam axis
  “EQUUS 81k M“ IR camera 3.0‐5.0μm wave length
  DN320 25cm length aluminium extensions – shift alonng beam
Limitation: sputtering at high power density high surface temperature  “cloud” between target and camera
COMSOL thermo-mechanical simulation
On-axis temperature vs time
Horizontal (-) and vertical (--) profiles after 200ms of NBI
NBI-1 (old IOS, 2017) 471kW 21keV (D) 200ms 3.30×14.63cm
NBI-1 (new IOS, 2019) 544kW 20keV (D) 100ms 4.12×7.80cm
NBI-2 (2021) 674kW
41keV (D) 95ms 2.34×5.40cm
DNBI (2024) 67kW 49keV 330ms
Thermography measurement on W-Target
NBI-1 “old” grids unacceptable divergency across slits (horizontal)
Lower horizontal beam size and divergency with new grids
Lower divergency and beam size for high energy NBI-2
Circular beam profile for axisymmetric IOS of DNBI
Collection optics: ×3 N-BK7  1 inch lenses, f:30mm on NBI-1/2 optical ports
View lines along (35-52 deg. to beam axis) or across (-12…+12 cm) beam
Bundle of 20 fused silica fibers, NA:0.22, core 365n arranged on 8mm in line
Czerny-Turner spectrometer – Acton Research SpectraPro-750, focal length 750 mm, aperture ratio f/9.7
AndorTM iXonEM+ Ultra DU 897 BV BI CCD Camera, 512×512 pixels, 16×16m pixels
Layout on NBI-2 (& NBI-1)
Measurement of Doppler-shifted (Balmer-, 656 nm) lines:
WL range 6519…6564 AO, 0.087 AO/pixel, instrumental function ~0.25AO
NBI & spectroscopy model
Spectroscopy measurement
Ballistic motion of beam particles form emitter (grids, G) 250mm = 2RG to point O
“Beamlet” axes to focal points (FX & F Y ~4m)
Velocity distribution from each point on emitter
f’(v) = jG × exp(– (’X/ X)2 – (’Y/ Y)2)
Velocity distribution at point (O) on view-line (OC)[5] : F’O(v) = G f’ dG
Projection of velocity distribution on view-line
F’’(vZ’’) =  Z’’ (nBG(z’’) × F’O(v) dz’’,
nBG(z’’) – background density along view-line
’X
’Y
Spectral distribution of Doppler shifted line on observation chord (OC)
F’’() = const × F’’(vZ’’),  = O (1 -(vZ’’/c)
Power density distribution PO(r) ≈ v E(v) F’O (v) dv
measurement NBI-2 49keV 900kW deuterium, along beam
model:  X& Y 0.8&0.3 deg.
Background density (nBG ) along beam axis – gas from neutralizer and calorimeter
jG =0 on IOS axis r15mm and cooling bar |z|5mm.
full energy
half energy
Eo/3
measurement NBI-2 52keV 1.1MW deuterium, along beam
Vertical divergence ( Y ) form measurement along beam
Larger line broadening for more perpendicular (to beam axis) view lines
magnet ON – only neutrals
 Y 0.28 deg.
magnet OFF – neutrals & ions
 Y 0.32 deg.
Line broadening vs observation angle (full energy fraction)
Vertical divergence from spectroscopy ~0.3 deg. is in good agreement with measurement on W-Target and calorimeter
Optimization of NBI divergency with spectroscopy measurement across beam
Minimal beam size at the same perveance for 3 beam energy fractions and for measurement on the calorimeter
IB < IOPT: profile broadening; IB > IOPT: redistribution form axis to 4-8cm
Beam halo, especially at low beam current
perveance scan
calorimeter
Eo Eo/2 Eo/3
IOPT
profiles vs IB
profile
redistribution
IB = IOPT
IB < IOPT
IB > IOPT
Summary
Measurement of NBI profiles with 3 systems implemented for NBIs on TCV:
  Calorimeter: power profile 10 points 1-D horizontal, max.500kJ (0.5s@1MW), fixed z, NBI shots on calorimeter
  W-target thermography: 2-D power distribution, different z, max. 100-200kJ, NBI disconnected from tokamak
  Beam spectroscopy: density of beam fractions, non-perturbative, > 2s@1.3MW, shots on calorimeter and in TCV