## Slide 2
01
Introduction
02
Design
03
Experiment
04
Conclusion

## Slide 3
The Principle of BNCT
Boron Injection: ¹⁰B drug selectively accumulates in tumor cells.
Neutron Beam: The tumor site is irradiated with a beam of epithermal neutrons generated by a neutron source..
Nuclear Reaction: ¹⁰B captures neutrons, causing ¹⁰B(n,α)⁷Li reaction.
The basic principle of BNCT is to utilize the capture reaction of boron-10 isotopes under neutron irradiation, selectively killing tumor cells by producing high-energy secondary particles of lithium and helium with high linear energy transfer (LET).
BNCT Process
Boron Neutron Capture Therapy(BNCT)
10B+n➡7Li+4He

## Slide 4
Accelerator-based BNCT
Breakthrough in high-current proton accelerator technology has shifted BNCT focus to accelerator-based systems, enabling medical commercialization.
Target
BSA
Accelerator
Neutron Irradiation System
Therapeutic System
Accelerator
Ions Source
RFQ
Target
BSA
Collimator
Patient
Patient
Li/Be Target
Iron、Aluminum、Heavy Water、Fluoride、Cement、W、Pd
Proton Beam
Cyclotron/DC High-Voltage/Linac Accelerator
Thermal Neutron
Neutron Beam
proton
Fast Neutron
Thermal Neutron

## Slide 5
Layout and Structure
ECR (Ion Source + RFQ (Radio Frequency Quadrupole) Accelerating Cavity + Solid-State Power Source:
Advantages: Low proton energy, high beam power, excellent stability, and long lifespan, making it particularly suitable for medical applications.
Low Proton Energy: Facilitates radiation protection, neutron moderation, reduces the size of the Beam Shaping Assembly (BSA), and shortens treatment intervals.
High Beam Power: Increases neutron yield, reducing treatment time.
35 keV IS
LEBT
2.787MeV RFQ
HEBT
TR1
TR2
TR5
Design parameters | Value
Output energy | 35 keV
Pulse intensity | 25 mA
Average beam current | ＞20 mA
RMS emittance at the exit | ＜0.2 π·mm·mrad
Duty cycle | 80%
Routine Frequency | 200Hz
The main design parameters for the ion source

## Slide 6
Challenges Encountered
Discharge: ECR ion source cannot enable 100% reliable discharging.
Alumina ceramics: the discharge chamber experienced instability due to the influence of the ceramic pieces, significantly affecting the plasma generation within the discharge chamber.
Bandwidth: the frequency dispersion of commercially available 2.45 GHz magnetrons is relatively large, with a nominal dispersion value of 50 MHz .

## Slide 7
Design Highlights
Low Energy: Reduce the difficulty of high-voltage insulation and enhance the high-voltage stability of the ion source, reducing or avoiding high-voltage discharging of the ion source.
Low SWR: Ensure easier discharging, guaranteeing the availability of the BNCT ECR ion source.
Axial Magnetic Field : Permanent magnets and adjustable pure iron for precision and flexibility.
One-click operation mode: Enable rapid startup of the ECR ion source.

## Slide 8
the S11 curve with cylindrical discharge chamber at the entrance of the ridged waveguide obtained through CST simulation
the S11 curve with square discharge chamber at the entrance of the ridged waveguide obtained through CST simulation
 | cylindrical discharge chamber | square discharge chamber
SWR | 1.6 | 1.06
bandwidth(MHz)* | 2.228 | 19.687
*Taking the frequency at which the S11 curve decreases to -3 dB as the start and end points for bandwidth
In the initial design, an ignition device was installed at the discharge chamber flange.
But it has never been used, indicating that the low SWR design of our ECR ion source enables 100% reliable discharging.
Low SWR Design
Advantages of Low SWR

## Slide 9
The magnetic field configuration provides a stable saddle-shaped magnetic field distribution within the discharge chamber, ensuring smooth and controllable electron motion, and enhancing plasma generation efficiency. The magnetic field is now provided by permanent magnets with a residual magnetization of 1.26 T.
Magnetic field Configuration
The magnetic field configuration obtained with the three permanent magnets
The magnetic field configuration obtained with the two permanent magnets and a pure iron

## Slide 10
Emittance and envelope at the entrance of first solenoid vs fraction of space–charge neutralization
Extraction system
The plasma electrode is set at a potential of 35 kV, the suppressor electrode at -2.2 kV, and the ground electrode at 0 V.
The RMS emittance was measured at a low duty cycle because the double-slit emittance measurement device is an interceptive instrument, where the first slit intercepts nearly the entire beam.
The normalized root-mean-square (RMS) emittance of the beam at the exit of first solenoid was 0.197 π·mm·mrad
Emissivity phase diagram at the exit of first solenoid

## Slide 11
Operating parameters of the BNCT02 ion source
Tuning and conditioning
The maximum pulsed current measured at 26.11 mA with a duty cycle of 75% for cylindrical discharge chamber

## Slide 12
Duty Cycle: 10% (32.56mA)
Duty Cycle: 20% (36.4mA)
Duty Cycle: 30% (38.8mA)
Duty Cycle: 50% (39.92mA)
Duty Cycle: 62.5% (41.52mA)
Duty Cycle: 75% (42.32mA)
Conditioning with Square Discharge Chamber

## Slide 13
Operation
Beam current of the ECR ion source with cylindrical discharge chamber at duty cycle of 60%
The clinical BNCT equipment has maintained with a targeted beam power of approximately 28 kW and an uptime rate of over 95%.
The ion source can provide beam current continuously for 24 hours, accumulating a total operating duration of 6,694 hours.
The power source has also been capable of continuous operation for 24 hours, with a total operating duration of 7,073 hours.
The neutron target has accumulated a total beam time of 5,160 milliampere-hours, with a single day’s longest beam time exceeding 110 milliampere-hours.
The neutron flux density of BNCT02 has reached the standards set by the IAEA for clinical BNCT treatments.

## Slide 14
Provide two different design of discharge chamber: cylindrical and square
It should be emphasized that the square discharge chamber exhibits higher stability and availability of the ion source for following reasons:
Lower SWR: Low SWR design ensures easier discharging, guaranteeing the availability of the BNCT ECR ion source.
Broader operational bandwidth: Discharge chamber with wider bandwidth can help ion source adapt to frequency fluctuations, which fully improves the stability of the ion source.
Currently, the clinical BNCT equipment installed at Dongguan People's Hospital has maintained stable and efficient operation over the past year, with a targeted beam power of approximately 28 kW and an uptime rate of over 95%.
The neutron flux density of BNCT02 has reached the standards set by the IAEA for clinical BNCT treatments.
Conclusion