## Slide 1: ARES @ DESY
Status update
VHEE conference 2025
Florian Burkart
DESY - MIL, injectors and linacs

## Slide 2: Since last VHEE 2023 … status update!
Very fruitful conference!
Unfortunately, we have a very hard budget crisis at DESY.
Reduction of ARES run-time in 2024 down to 3 month.
Operation team reduced from 5 to 2 Postdocs.
Heavy impact on availability.
2025 looks much better but still reduced availability.

## Slide 3: Deutsche Elektronen Synchrotron (DESY)
Campus in Hamburg Bahrenfeld

## Slide 4: DESYs accelerator landscape
Campus
FLASH
PETRA III
DESY II
EuXFEL
FLASH II
SINBAD
ARES Linac
KALDERA

## Slide 5: ARES, A Precision Tool for Accelerator Science, Technology and Application Developments
            ARES goals:
Generate and characterize ultrashort e- bunches (fs to sub-fs) with high stability for applications related to accelerator R&D.
Now strong focus on medical applications,VHEE & FLASH, with dedicated experimental area.
Parameter | Target | Status
Charge | 0.05 - 200 pC | 0.003 - 200 pC
Momentum | 50 - 150 MeV/c | 20 - 160 MeV/c
Momentum Spread | 10-4 | 10-4 (res. limited)
Rep rate | 10 Hz | 1 – 50 Hz
Bunch Length (rms) | Sub-fs to ~10 fs | ≈ 2.9 fs
Dose / pulse |  | 2 – 5 Gy

## Slide 6: The 3 pillars for medical research
Strong collaboration to develop and test fast & precise dosimetry.
Dosimetry
VHEE & FLASH cell irradiation
Medical Imaging
Starting with cell irradiations,
VHEE / FLASH beam parameters with medical partners.
Medical imaging – the electron CT,
DESY internal collaboration

## Slide 7: Infrastructure for medical experiments

## Slide 8: Three Experimental Areas for Users
Full Flexibility
In UHV (~7e-10 mbar), highest precision,
best beam control.
In air, highest flexibility.
In low pressure vacuum, detector test stand.

## Slide 9: S2 Bio lab
Available for medical experiments at ARES.
provides all the necessary equipment to cultivate, store, manipulate and analyse biological agents
BSL2 safety workbench equipped with a mini centrifuge, vortexer, vacuum pump,…

## Slide 10: Dosimetry

## Slide 11: The challenge of „the dose“…
How to translate electrons to dose?
Medical side ~ translator ~ accelerator side
What is the „right“ dose?
  Integrated dose
  Dose rate
  Dose rate/pulse
  How do we measure dose (fast!)

## Slide 12: Dosimetry development with the Wollongong university (Australia)
Collaboration with UKE & Woolongong university within the framework of HiACTS (Helmholtz Innovation Platform for Accelerator-based Technologies and Solutions) for the development of an online dosimetry monitor (Moskin)
First beam times to compare Moskin with several other dosimeters (analysis of the results is ongoing). Next beam time in the coming months.
courtesy: James Cayley, Elisabetta Gargioni

## Slide 13
Testing and benchmarking different detector systems (NitroFLASH, diamond, films, PANDORA, etc.)
Other detector R&D activities
Dosimetry and radiation detector tests from the DESY radiation protection group
See talks from Oliver Stein
and Florian Beißer

## Slide 14: Depth-dose curves in (solid) water
Setting up tools and simulations
Energy = 153 MeV
Beam: approx. Gaussian
FWHM = approx. 2.8 mm (both in x and y) @ phantom surface
About 30 pC per pulse
Dose measured with EBT3-XD films: approx. 0.55 Gy/pulse
Monte Carlo simulation with experimentally-determined beam divergence
Courtesy: Joel Senador, UKE

## Slide 15: The results: Beam / Dose structure @ ARES
Single pulses with up to 50 Hz rep. rate
1-5 Gy
1 / 0.1 / 0.02 sec
Single digit ps
O (TGy/s)
O (50 – 250 Gy per second)
Current Limitations:
Only 250 pC per pulse
 Small spot size in the order of 10 mm2

## Slide 16: Cell irradiations

## Slide 17: ARES collaboration with UKE
Prof. Dr. K. Rothkamm
+ Dr. N. Struve, Dr. E. Gargioni
+ J. Senador et al.
Collaboration meeting with colleagues from UKE, DESY, Uni Lund and Uni Wollongong
Very high energy electrons (VHEE) & FLASH treatment
Studies at ARES for Dosimetry
Third party funding received in collaboration with Uni Wollongong (Australia) for online dosimetry development.
Studies concerning treatment of medulloblastoma
Master thesis of Joel Senador (now PhD)

## Slide 18: How to hit the cells with small beams?
Preparation in the lab and transport to ARES is not an issue.
Bringing the cells in the tunnel is also not a problem (special access mode to keep most of the accelerator parameters).
Part 1
Eppendorf tubes are not Eppendorf tubes.
Different holder systems from DESY and UKE
Combination with online dosimetry.

## Slide 19: How to hit the cells?
ARES is limited in the amount of charge we can produce / transport.
Small spot size, where we can achieve FLASH like beam parameters.
Cells need to be prepared very carefully.
We have to rely on dosimetry!
Still the issue remains: are we sure that we hit the cells with the right dose?
Set-up of different irradiation patterns: Spot (FLASH like parameter) vs. Snake pattern (VHEE parameters, uniform distribution) irradiation
Part 2

## Slide 20: Developing routines and first results
Analysis ongoing, publication in preparation
Experimental setup, cells in Eppendorf tubes
Dose analysis after irradiation
Survival rate of Med8a cells
Last 2 weeks: increase charge/dose, statistics for cell irradiations.
Courtesy: H. Wanstall, R. Jones
First in vitro measurement of VHEE relative biological effectiveness (RBE) in lung and prostate cancer cells using the ARES linac at DESY

## Slide 21: Regular beam times with UKE, tumor cell irradiations
Working on statistics
Irradiation of different cell lines with 2 – 10 Gy
Beam studies (bunch length, optimization of charge density)
Preparation for publications

## Slide 22: Studies with different types of samples
Studies for tumor treatment with high energetic electrons and FLASH parameters.
Mosfet detectors from Uni Wollongong as online dosimetry.
Animal and skull phantoms and living tumor cells.
Preparation for egg experiments.

## Slide 23: Medical imaging

## Slide 24: eCT - Medical imaging and treatment with one accelerator
DESY cross-divisional research project
First proof-of-principle tomography successful
Next steps:
Tests with more phantoms from University Hospital UKE
What can we detect?
Improve data acquisition system and reduce measurement time.
ElectronCT – an inmaging technique using very-high energy electrons, P. Schuetze et al.,
Front. Phys., Sec. Medical Physics and Imaging, Volume 12 - 2024 | https://doi.org/10.3389/fphy.2024.1454854
Timepix3 detector
Mouse phantom
courtesy: Paul Schuetze

## Slide 25: Summary

## Slide 26: ARES - summary
ARES – A state of the art normal conducting linac with highest stability and reproducibility.
Testbed for internal and external Accelerator R&D – full flexibility with three different experimental stations
Rich portfolio of completed, ongoing, and planned R&D activities
Unique opportunities for medical applications
Towards one accelerator for imaging and treatment with VHEE and FLASH radiotherapy beam parameters.
Platform for dosimetry detectors developments.
Ramping up activities again in the second half 2025.

## Slide 27: Thank you!
Florian Burkart
DESY - MIL (injectors and linacs)
florian.burkart@desy.de
+49 40 8998 3039
Notkestr. 85, 22607 Hamburg