## Slide 1: Dosimetry Codes of Practice Are they applicable to VHEE RT?
Hugo Palmans
National Physical Laboratory, Teddington, UK and MedAustron, Wiener Neustadt, Austria

## Slide 3: TRS-398 Rev.1:

## Slide 4: IAEA TRS-398 Rev.1: Framework/ Simple absorbed dose formalism
But we have
with :
If = 60Co:
SI definition of quantity:
* Absorbed dose (to water) in broad beam
Primary standards:
* Calorimeters, ICs, Fricke
Secondary standard:
* Ionization chamber
Relative dosimetry in water:
* Output factors
* Profiles 1-2-3D
* TPS
Quantity at user:
* Reference absorbed dose
* Beam monitor calibration
TRACEABILITY

## Slide 5: TRS-398 Rev.1: recommendations on equipment

## Slide 6: TRS-398 Rev.1: recommendations on beam quality Q

## Slide 7: For which field size and how to measure R50 in VHEE beams?
(Böhlen et al 2021 Med Phys 48:3958)

## Slide 8: How to measure R50 in focused VHEE beams?
(Fan et al 2023 – IPAC23)

## Slide 9: TRS-398 Rev.1: reference conditions for reference dosimetry:

## Slide 10: Where does this come from?

## Slide 11: Where does this come from?

## Slide 12: Where does this come from?

## Slide 13: VHEE data compared to Burns
But: that Burns value would be for a reference depth of 30 cm…

## Slide 14: TRS-398 Rev.1: beam quality correction factors kQ
()

## Slide 15: Beam quality correction factors kQ

## Slide 16: TRS-398 provides recommendations on the measurement of depth-dose distributions, profiles and output factors

## Slide 17: TRS-398 provides recommendations on some other things
Correction of the raw ionization chamber reading for influence quantities
Use of plastic phantoms
Cross calibration of field ionization chambers
(Petersson et al 2017 Med Phys 44:1157)

## Slide 18: TRS-398 provides recommendations on some other things
Correction of the raw ionization chamber reading for influence quantities
Use of plastic phantoms
Cross calibration of field ionization chambers

## Slide 19: TRS-398 provides example/typical uncertainty budget

## Slide 20: So, when new modalities come that deviate from those that led to TRS-398 Rev.1, one has to consider:
Can the same equipment be used?
Are the reference conditions still valid or can they even be achieved?
Is the beam quality specification still relevant and can it be measured in the same way?
Can the same kQ values be used?
Are the recommendations on relative dosimetry still valid?
Is the typical uncertainty budget still valid or are any components much larger?

## Slide 21: UHDR proton RT
Can the same equipment be used? √ but some improvements and alternatives
Are the reference conditions still valid or can they even be achieved? √
Is the beam quality specification still relevant and can it be measured in the same way? √
Can the same kQ values be used? √
Are the recommendations on relative dosimetry still valid? √
Is the typical uncertainty budget still valid or are any components much larger? √

## Slide 22: (UHDR) VHEE RT
Can the same equipment be used? √ but some improvements and alternatives for UHDR
Are the reference conditions still valid or can they even be achieved? √ to some extent but relevance should be questioned
Is the beam quality specification still relevant and can it be measured in the same way? √ probably yes
Can the same kQ values be used? X maybe with different model, but more data needed
Are the recommendations on relative dosimetry still valid? √ maybe, but more data and extension needed
Is the typical uncertainty budget still valid or are any components much larger? √ but some larger contributions

## Slide 23: UHDP/R solutions
Ultra-thin PP IC FLASH diamond
Gomez et al 2022, Med Phys 49:4705
Marinelli et al 2022, Med Phys 49:1902

## Slide 24: Calorimeters as reference instruments in the clinic for UHDP/R

## Slide 25: SFRT VHEE beams
Can the same equipment be used? √ to some extent
Are the reference conditions still valid or can they even be achieved? X
Is the beam quality specification still relevant and can it be measured in the same way? X
Can the same kQ values be used? √ maybe but more data needed
Are the recommendations on relative dosimetry still valid? X
Is the typical uncertainty budget still valid or are any components much larger? X

## Slide 26: SFRT - average dose measurement
Cotterill et al 2024, submitted to Cancers

## Slide 27: SFRT - peak or valley - detector orientation, e.g. microDiamond
Small field mini beam

## Slide 28: Formalism
IAEA TRS-398:
IAEA TRS-483:
(small field)
MRI/linac:
(de Pooter et al PMB 66:05TR02)
UHDP/R beams:
VHEE beams:
SFRT beams:

## Slide 29: In conclusion
No recommendations for VHEE beams in existing COPs for reference dosimetry in external beam radiotherapy
Framework and formalism can still be used, but may have to be expanded for SFRT beams
Relevance of beam quality and reference depth for HE electrons is questionable for VHEE
More systematic determination of stopping power ratios and beam quality correction factors are needed
For UHDR VHEE alternative instrumentation is required
For SFRT reference conditions and appropriate quantities need to be better defined and positioning is a larger issue