## Slide 1: CLEAR, operational status and VHEE activities
4th International Workshop on Very High Energy Electron Radiotherapy
Daresbury, UK
Giacomo Tangari - giacomo.tangari@cern.ch

## Slide 2: CLEAR, operational status and VHEE activities
4th International Workshop on Very High Energy Electron Radiotherapy
Daresbury, UK
Giacomo Tangari - giacomo.tangari@cern.ch

## Slide 3: What is CLEAR?
CERN Linear Electron Accelerator for Research
Versatile linear accelerator facility, supports R&D in accelerator technologies and applications.
Provides electron bunches with energies up to ~200 MeV, flexible bunch structure and timing.
Key applications:
  Accelerator component testing (e.g., diagnostics, kickers, RF systems)
  High-gradient structure experiments
  Plasma wakefield acceleration studies
  Machine Learning
  Medical beam applications

## Slide 4: CLEAR 2nd Beamline

## Slide 5: Operational Tools
CLEAR DAQ
  Data acquisition system for automated, synchronized collection of beamline measurements.
CLEAR Online Dosimetry
  Real-time dose measurement and feedback for radiation-sensitive applications.
CLEAR Robot
  In-house developed robot, leveraged during irradiations to move holders from the storage box to the beam trajectory.
Scan Worker
  Framework to automate multi-parameter scans (e.g., laser, RF settings) and coordinate data acquisition.
CLEAR Simulations
  Accompanying simulation framework to compare with experimental results.

## Slide 6: Film Dosimetry
Use of Gafchromic films placed in direct contact with the biological samples during irradiation.
Films provide the effective experimental dose, independent of real-time predictions.
They also act as a “photograph” of the beam, revealing key information:
  Beam profile and uniformity → ensures homogeneous irradiation of the sample area.
  Delivered charge → correlation between accelerator settings and actual dose.
  Additional effects → detection of contributions such as dark current and leakage, which alter the intended dose distribution.
Up to 5 Gy of extra dose for 2m30s exposure with laser shutters closed
To be checked

## Slide 7: Dose Perturbations
Observed Secondary Contributions can lead to non-negligible impact on the experiments.
Dark Current
  Defined as unwanted electron emission from RF structures without laser injection.
  Increases linearly with irradiation time.
  Produces a “dose tail”.
Leakage
  Acts as a background dose, always present in the beamline.
  Reduces spatial resolution and beam concentration on samples.
Next steps: investigate strategies to mitigate dark current and leakage.
DC w/ screen | DC w/o screen | Leakage
1.40 pC | 2.20 pC | 6.20 pC
1.23 pC | 2.18 pC | 6.36 pC
1.30 pC | 2.15 pC | 6.90 pC
Average charge values measured at the end of the line, obtained from 1000 pulses (10 Hz repetition rate over 100 seconds).

## Slide 8: Shorter Bunches
Possible Strategy: increase the charge density of the beam.
Achieved by shortening the bunch length.
Check the instantaneous dose rate and the subsequent effect based on the beam delivery:
  In Conventional high charge density may increase complex biological damage and decrease sparing effect;
  In FLASH it could trigger a temporal-biological sparing effect.
First measurement campaigns are confirming the procedure and, thus, the charge density increase.
But is it the right path?

## Slide 9: References and Publications
Papers, used as a reference in this presentation, regarding CLEAR can be found at clear.cern/content/publications.
Some of the presented work is preliminary and has not been published yet.
Many thanks to the whole CLEAR team and the various users.

## Slide 10: Questions? Ideas?
Thank You!
clear.cern.ch

## Slide 11: Backup Slides

## Slide 12: CLEAR Parameters
Beam parameter | Range
Energy | 30 – 230 MeV
Energy Spread | < 0.2 % rms (< 1 MeV FWHM)
Bunch Length | 0.1 ps – 10 ps rms
Bunch Charge | 5 pC – 3 nC
Number of bunches per pulse | 1 to ~150
Maximum total pulse charge | 80 nC
Normalized emittances | 3 mm to 30 mm (bunch charge dependent)
Repetition rate | 0.8 to 10 Hz
Bunch spacing | 1.5 GHz or 3 GHz

## Slide 13: CLEAR Line
CLEAR New Beamline

## Slide 14: 2nd Beamline Photos