## Slide 1
Supersonic gas curtain-based in-vivo transverse beam profile monitoring for medical accelerators
Dr Narender Kumar
Research Coordinator

## Slide 2
Background and R&D
Narender.Kumar@Cockcroft.ac.uk
MCP
Phosphor screen
Camera
Ionization
Ions extraction
Visible light
Primary Beam

## Slide 3
Gas jet shaping
Background and R&D
Argon
Nitrogen
Narender.Kumar@Cockcroft.ac.uk

## Slide 4
Supersonic Gas Curtain based real-time Ionization Profile Monitor
Status |  | Pressure (mbar) |  |  | 
Gas Jet | Beam | Nozzle Chamber | Skimmer Chamber (b/w 1st - 2nd) | Skimmer Chamber (b/w 2nd - 3rd) | Interaction Chamber
Off | Off | 3.98 × 10-08 | 7.10 × 10-09 | 3.00 × 10-09 | 1.73 × 10-08
On | Off | 4.84 × 10-03 | 5.60 × 10-06 | 4.60 × 10-07 | 2.13 × 10-08
On | On | 4.93 × 10-03 | 5.60 × 10-06 | 4.60 × 10-07 | 2.16 × 10-08
Narender.Kumar@Cockcroft.ac.uk

## Slide 5
Proof-of-Concept Measurements
Pelletron Accelerator facility, Dalton Cumbrian Facility (DCF), United Kingdom |  |  |  |  | 
Beam Species | Charge state | Beam Energy (MeV) | Average Beam Current (nA) | Beam Size | Gas Curtain Species
Proton | +1 | 4-8 | 10-100 | Up to a few mm’s | Argon, Nitrogen
Carbon | +2, +3, +4, +5 | 12-24 | 2-100 | Up to a few mm’s | Argon, Nitrogen
MC40 Cyclotron Facility, University of Birmingham (UoB), United Kingdom |  |  |  |  | 
Beam Species | Charge state | Beam Energy (MeV) | Average Beam Current (nA) | Beam Size | Gas Curtain Species
Proton | +1 | 10.8- 28 | 2-72 | Up to a few mm’s | Argon, Nitrogen

## Slide 6
Proof-of-Concept Measurements
What is ideal for our device
Stable beam
Good vacuum (<10-8 mbar), Low noise
What is not ideal for our device
poor vacuum (10-7), strong background signal
Difficult to isolate noise
Narender.Kumar@Cockcroft.ac.uk

## Slide 7
Evaluation of Detector Performance
Besides shape, the only quantitative number we get is total Counts (C) on pixels.
Input beam
flux distribution
Imaged beam
count distribution
Assuming no distortion of profile during extraction
Probability of ionisation (function of energy)
Detector specific factors
Scaling factor
Based on detector performance, a term was introduced Detection Limit (D)
Narender.Kumar@Cockcroft.ac.uk

## Slide 8
Performance of the Detector
Variation of gas
D values for 100nA beam current from DCF campaign
C variation for 28 MeV P-beam at different beam currents
Detection limit (D) variation with gas species (Ionization cross section) and gas density dependence
Detection limit (D) variation over energy (Bethe-Bloch Equation for energy loss)
Detection limit (D) stability over different integration time (dependent only on detector configuration)
Counts (C) linearity with Beam current (potential for ultra high dose rates (UHDR))
Narender.Kumar@Cockcroft.ac.uk

## Slide 9
Summary of POC Measurements
Pelletron Accelerator facility, Dalton Cumbrian Facility (DCF), United Kingdom |  |  |  |  |  |  |  | 
 | P+ (12MeV) |  | P+ (15MeV) |  | P+ (127MeV) | C6+ (33.4MeV/u) |  | 
Instantaneous dose rate | 1.00 x 109 |  | 1.80 x 109 |  | 1.56 x 1010 | 9.70 x 108 |  |  LhARA design baseline
pulse length (ns) | 7 |  | 7 |  | 41.5 | 75 |  | 
Charge per laser shot (pC) | 100 |  | 100 |  | 100 | 100 |  | 
ions per bunch (no.) | 6.24 x 108 |  | 6.24 x 108 |  | 6.24 x 108 | 1.04 x 108 |  | Calculated
pencil 1 mm (ions/sq.mm/bunch) | 7.95 x 108 |  | 7.95 x 108 |  | 7.95E x 108 | 1.32 x 108 |  |  Fluence
Pencil: 1 mm | Ar | N2 | Ar | N2 |  | Ar | N2 |  Number of bunches required
 | 42 | 52 | 52 | 63 | x | 8 | 20 | 
Based on POC Measurements at UoB |  |  |  |  |  |  |  | 
 | P+ (12MeV) |  | P+ (15MeV) |  | P+ (127MeV) | C6+ (33.4MeV/u) |  | 
Pencil: 1 mm | 0.17 | 0.21 | 0.21 | 0.26 | x | 0.03 | 0.08 | Number of bunches required
Narender.Kumar@Cockcroft.ac.uk

## Slide 10
Outlook
More compact and flexible design for easier system integration.
Optimization of gas jet density at interaction point to meet vacuum requirements.
More experiments to identify gases with higher ionization cross-sections
Upgrade to extraction system for enhanced detection limit.
Size Reduction: 25% along beam, 50% lateral to beam
Extraction unit
Detector
Chamber and support structure
Gain Improvement |  |  |  | 
 | Pessimistic | Optimistic | Actual tested at UoB | 
QE P43 | 1.2 | 1.4 | 1.3 |  Manufacturers datasheet
MCP Double chevron | 100 | 1000 | 80 |  Operation limited by poor vacuum
Effective ions extraction | 1 | 1-4 | - |  New Design being manufactured
Curtain thickness | 2 | 5 | - | Unchanged
Curtain density | 2 | 5 | 3 |  ZNS optimization
 | 480 | 49700 | 240 | 
Narender.Kumar@Cockcroft.ac.uk

## Slide 11
Outlook
Narender.Kumar@Cockcroft.ac.uk

## Slide 12
Optical Air Fluorescence Monitor
Narender.Kumar@Cockcroft.ac.uk
Camera
Image Intensifier
25:7 imaging optics
Band pass Filter wheel
Apochromatic Triplet F= 600
Aperture
Depth of field
Beam
Beam exit
AFM installed at MC40 cyclotron (August 2024)

## Slide 13
Conclusions
Supersonic gas jet ionization profile monitor has successfully demonstrated capability for 2D transverse beam profile measurements with beams close to the requirements for ion beam cancer therapy.
The design is updated to improve the detection limit and easier integration possibility
POC measurements were carried out for OAFM successfully. More measurements (to be planned)
Last 18 months: 3 conference talks, 7 posters, 8 conference proceedings, 3 journal articles (under review), 3 journal articles (to be submitted soon).
Special thanks to the JetDose team: M. Patel, W. Butcher, F. Thesni M. P. and C. P. Welsch
Thanks to Dr. Andrew Smith (DCF) and Dr. Tony Price (UoB) for the assistance with the POC measurements
Narender.Kumar@Cockcroft.ac.uk