## Slide 1: Recent RF test results of Nb/Cu full-seamless 1.3 GHz cavities
Hayato Araki
High Energy Accelerator Research Organization (KEK)

## Slide 2: KEK/CERN Collaboration for Nb/Cu R&D
  Nb/Cu cavity development from LEP
  HiPIMS coating facility
  Copper tube hydroforming
  FCC
  Bulk Nb cavity development from TRISTAN
  1.3 GHz processing facility
  Bulk Nb hydroforming
  ILC
CERN
KEK

## Slide 3: 1.3 GHz cavity fabrication
Measurement
Coating
Hydroforming
Simulation

## Slide 4: Cavities overview
C1, C1b
The first series of full-seamless
KEK14, KEK15
Brazed and welded with CF flanges
C1c, C1d
Hydroforming + Plasma EP
C3, C3b
Mechanical polishing + Hydroforming

## Slide 5: Nb thin film coating
Process parameters
  EP: 40 minutes, target 12 um removed material.
  HPR: 100 bars, 3 hours, cavity dried under vacuum at 50°C
  Bakeout: 48h, 200°C, base pressure ~ 2×10-10 mbar
  Coating: HiPIMS
    Cut-offs: 350W, 2.5×10-2 mbar Kr, 6 positions/cut-off, 150°C, unbiased
    Cell: 2 kW, 2.5×10-3 mbar Kr, 6 hours, 150°C, -75V bias voltage
    Targeted thickness at the equator: 6 um
https://doi.org/10.1016/j.surfcoat.2022.128306

## Slide 6: The first series
C1 and C1b (2024-03)
  Our target: Q0 > 4×108 at 12 MV/m, 4.2 K
    Acc. gradient: Cleared
    Quality factor: Just short of the target
  1.85 K
    High-Q in low field and Q-drop
    Max Eacc: 15.5 MV/m

## Slide 7: Measurements at CERN
KEK14, KEK15, C1 and C1b (2024-06 to 2025-07)
  Consistent results came out from the same cavity measured in both KEK and CERN
  Improved cool-down method
    Spatial temperature difference ΔT at Tc affects surface resistance.
    Covering the cavity with a cylindrical shield worked well.
KEK14: measured w/o the shield for small ΔT cooling

## Slide 8: Thermoelectric current
Large ΔT cooling (without shield)
Small ΔT cooling (with shield)
Small ΔT cooling is essential for bilayer structure cavities.
A simple “shield” is effective for vertical cryostat measurement.

## Slide 9: Substrate polishing
C1d (2025-08)
  Plasma EP (~200 um) after HF
  Max. Eacc was higher (18 MV/m) than C1/C1b
  Q-drop has occurred
C3 and C3b (2025-06 to 10)
  Mechanical polishing + light EP before HF
  Two cavities received the same treatment
    C3b kept Q0 > 1×1010 until 17 MV/m (1.85 K)
    C3 had Q-drop and steep Q-slope  A locally oxidized spot on the substrate

## Slide 10: Surface condition comparison
 | C1, C1b | C1d | C3, C3b
Treatment | Baseline | Plasma EP ~200 um after HF | Mechanical polishing before HF
Surface |  |  | 
Roughness (Ra) | 3.3 um | 0.69 um | 2.7 um
Q0 at 10 MV/m | C1: 3 ×109 C1b: 3 ×109 | C1d: 4 ×109 | C3: 1 ×109 C3b: 3 ×1010

## Slide 11: Additional substrate polishing
C3 after 2nd MP (2026-09)
  Cu substrate had 2nd mechanical (buff) polishing followed by Nb re-coating.
  Q in low field was lower than before, but Q-drop did not occur.
  Quench field 21.5 MV/m is the highest in our Nb/Cu cavities.
  Target cleared: Q0 > 4×108 at 12 MV/m, 4.2 K

## Slide 12: Discussion
Recovery from Q-drop
  Is “a local defect on substrate” the cause of Q-drop?
  Which stage of the manufacturing process is the cause of the defect?
    In HF cavities, Q-drop occur more frequently than in other seamless cavities.
    Copper material? Tube drawing? Hydrant? Coating preparation?
  Is it possible to recover from Q-drop using other polishing techniques?
Nature of Nb film
  Anti-Q-slope? Dirty?
    In low field (< 2 MV/m), the residual resistance decreases as Eacc increases.
    BCS resistance decreases up to 12 MV/m.
    MFP of HiPIMS Nb film is shorter than that of ILC-treated bulk niobium.
Kristof Brunner, TFSRF2024
https://indico.cern.ch/event/1376902/contributions/6111630/

## Slide 13: Summary
Common features in our Nb/Cu cavities
  “Q-drop” frequently occurs.
  Spatial temperature gradient at Tc has strong correlation to the residual resistance.
Substrate polishing
  At present, no relationship between surface roughness and RF performance can be confirmed.
    One cavity kept Q0 > 1×1010 until 17 MV/m with orange-peeled surface.
  “Q-drop” occurred in one cavity with locally oxidized spot on the substrate; however, this was recovered by mechanically polishing the substrate.
    Max Eacc reached 21.5 MV/m at 1.85 K, the highest in our cavities.
    Our target: Q0 > 4×108 at 12 MV/m, 4.2 K was cleared.
Acknowledgements