## Slide 1: Advancements in MgB2-Coated Cu Superconducting RF Cavities
Ke Chen, Xiaojun Xu, Dmitriy Dikin, and Xiaoxing Xi
Temple University, Philadelphia, Pennsylvania, USA
Eric Lechner
Thomas Jefferson National Accelerator Facility, Newport News, Virginia, USA
Leila Kasaei
Rutgers University, New Brunswick, New Jersey, USA

## Slide 2
 | Nb | MgB2
 | 9.2 | 40
 | 5 | 0.1
 | 1.5 | 7 (σ), 2(π)
 | 0.20 | possibly 0.80
 | 0.17 | <0.1
 | Nb | MgB2
 | 9.2 | 40
 | 5 | 0.1
 | 1.5 | 7 (σ), 2(π)
 | 0.20 | possibly 0.80
 | 0.17 | <0.1
RF surface resistance depends on energy gap and residual resistivity. Larger gap and lower resistivity indicate potential low RF loss than in Nb.
Field gradient is ultimately limited by thermodynamic critical field. For MgB2, Bc(0) could be as high as 800 mT, vs. 200 mT for Nb.
Lower critical field Bc1(0) may be an obstacle.
Goal: Coat 1.3 GHz Cu cavities with MgB2 films for particle accelerators
Potential to operate at 10-20 K with similar or higher energy gradient than Nb.
MgB2 vs Nb
Charged
Particles

## Slide 3
Mg: thermally evaporation of Mg pellets (physical vapor) creates high Mg vapor pressure
B: thermal decomposition of B2H6 (chemical vapor)
H2 atmosphere prevents oxidization of Mg
Liquid phase of Mg-Cu alloy at > 485˚C helps to reduce the growth temperature to about 500˚C - first depositing Mg before depositing MgB2
Growth of MgB2 Films by Hybrid Physical-Chemical Vapor Deposition (HPCVD)
Sample
heater
Mg heater
Sample
(up to 4.5” dia.)
Mg shots
B2H6 + H2
Pump
7.9” dia.

## Slide 4
HPCVD deposition system for large samples up to 4.5” in diameter
Deposition system
Inside the reactor chamber
10 cm
MgB2-coated Cu parts for axion detection
MgB2-coated Cu disks and ring for dielectric resonator

## Slide 5
4”-diameter Cu disk pre-cut on the back for easy cutting after deposition
25 1cm x 1cm samples were measured by AC susceptibility to determine Tc0
Tc0’s vary between 36 to 37 K, indicating uniform film quality
4”-diameter Cu disk coating and characterization by AC Susceptibility

## Slide 6
Superconducting Transition of MgB2 Films on Various Metals
Provides future options of interlayer to block interdiffusion of Cu into MgB2
MgB2 films on brass exhibit increased Tc0 up to 41 K, dropped to 39 K after thermal cycling to room temperature. Only reports of Zn-doped MgB2 with decreased or little changed Tc. Metastable phases?

## Slide 7
5-sample AC susceptibility measurement
5 samples cooled down to cryogenic temperatures simultaneously
AC susceptibility measurement conducted at 0.5 Hz to reduce signals from 1mm-thick Cu substrates
Each sample measured at temperatures from 5 to 50 K at rate of 0.5 K/min
10 samples can be measured per day

## Slide 8: Coating of mockup cavities and witness samples
Half-cell mockup
Cu cavity
(from Jefferson Lab)
2 clam-shell heaters formed by coaxial heating elements fixed on stainless-steel thermal shield
5 heating zones:
1&5 flanges
2&4 neck tubes
3 center
with separated heaters and thermocouples
1
2
3
4
5
Step-motor driven linear manipulator to provide vertical movement of up to 500 mm
IEEE Trans. Appl. Supercond. 36, 3500505 (2026) https://doi.org/10.1109/TASC.2026.3653724

## Slide 9
Chilled water in
Chilled water out
H2+B2H6
Mg shots
H2+B2H6
Gas feeding tube Up/Down moveable
Movable Mg and B sources & growth conditions
Growth conditions:
Gas: H2 (500sccm), B2H6(5%)+ H2(95%) mixture (75sccm). Total pressure: 5 Torr
Mg furnace 635℃, cavity heaters 465℃
Steps:
  Cavity heaters and Mg furnace heat up to target temperatures;
  Mg furnace heater on, B2H6 off, sources move through from top to bottom of the cavity in 20 min, slower at the center of the cavity;
  Mg furnace heater off, B2H6 on, sources move through from top to bottom of the cavity in 20 min again;
  Both Mg furnace heater and B2H6 gas are off, keep cavity heaters temperature at 465℃ to anneal for 20 min.
  All heaters and gases are off to cool in H2 till room temperature before cavity being taken out.
Mg furnace

## Slide 10
Sample# | Location | Tc0 (K) | Tc (K) (10~90%)
A | Top Tube | 37.5 | 2.5
B | Top Iris | 37.1 | 0.8
C | Top Middle | 37.1 | 1.2
D | Top Equator | 37.2 | 2.7
E | Bottom Equator | 37.1 | 3.1
F | Bottom Middle | 37.3 | 3.9
G | Bottom Iris | 37.6 | 3.3
H | Bottom Tube | 37.1 | 2.8
Sample# | Location | Tc0 (K) | 
A | Top Tube | 37.5 | 2.5
B | Top Iris | 37.1 | 0.8
C | Top Middle | 37.1 | 1.2
D | Top Equator | 37.2 | 2.7
E | Bottom Equator | 37.1 | 3.1
F | Bottom Middle | 37.3 | 3.9
G | Bottom Iris | 37.6 | 3.3
H | Bottom Tube | 37.1 | 2.8
AC susceptibility measurement of witness samples
Eight 10mm x 10mm x 1mm Cu substrates mounted on various locations of the inner wall of the cavity
All Tc’s are 37.1~37.6 K, close to highest reported Tc (38K) of MgB2 on Cu
Tc’s are 0.8~3.9 K, reasonably small, indicating full sample transit to superconducting below 33 K

## Slide 11
RF Measurement via Dielectric Resonator Technique
Sapphire rod: Φ0.315” X 0.398”
MgB2-coated Cu cavity inner diameter Φ1.25”
Measured in a cryocooler down to 6 K

## Slide 12
Loaded Q-factors and resonate frequencies: MgB2 vs Nb

## Slide 13
The measured Q (loaded) is the total loss:
1/Q = 1/QSC_film + 1/Qother
1/Qother is estimated at the level marked by the horizontal line in the graph.
RNb(9.2K) ~ 5 mΩ, RNb(7K) ~ 0.01 mΩ = 0.02 RNb(9.2K)
Assuming that 1/Qother are the same for both measurements, RMgB2(6K) ~ 0.35 RNb(9.2K) = 1.8 mΩ
Preliminary Analysis of Surface Resistance of MgB2 /Cu at 10.8 GHz
Possible culprit: Relatively large residual resistance of MgB2 at low temperatures
Good news: Better than Cu: Rs about 1/4 of Cu (RRR=400), which was measured to be 7.2mΩ for 4~20 K at 11.4 GHz, (Cahill et al., Proceedings of IPAC2016.) Translating to ~1/300 of Cu for 1.3 GHz.
Pogrebnyakov et al., APL 82,4320 (2003)
HPCVD
MgB2 film
RRR ~ 30

## Slide 14
Mg
Cu
Cu substrate
~2um
Mg-Cu alloy
~4um MgB2
(cross-section)
MgB2
(tilted surface)
SEM
B | C | O | Mg | Cu | Total | Location
72.99 | 7.05 | 0.92 | 13.49 | 5.55 | 100 | Spectrum 7
56.83 | 2.21 | 0.9 | 27.27 | 12.79 | 100 | Spectrum 8
Top View
Cross-section View
XRD
Both Mg2Cu and MgCu2 are present in the sample. These Mg-Cu alloys may be crucial for the growth of MgB2 films on Cu as it provides Mg needed.
4 um thick MgB2 film is likely sufficient for the SRF cavity applications.

## Slide 15: Rutherford Backscattering Spectrometry (RBS) for MgB2 film on Cu substrate
Fitting notes:
Using MultiSIMNRA program to fit with fixed B:Mg=2:1 for the first ~12 layers, then keep B=0 for 2 layers for Mg-Cu alloy.
Results:
Cu exists in the MgB2 films, possibly in a separated phase.
The concentration of Cu is not zero at the surface.
The thickness of the MgB2 film is around 2um.
The Mg-Cu alloy beneath MgB2 is likely only MgCu2.
Sample
θ = 163o
5 cm
β = 17o
2.4 MeV 4He++
α = 0
Detector
Past results show Cu inclusion inside MgB2 films. [Withanage et al., Supercond. Sci. Technol. 30 (2017) 045001]
New FIB/SEM/Laser triple beam system may help profile MgB2 films

## Slide 16: Summary
MgB2 films coated by HPCVD on Cu studied for SRF cavities.
High Tc up to 39 K obtained on Cu, brass, stainless steel, Mo, Hastelloy, Ni.
Uniform films with high Tc obtained on 4” dia. Cu disc.
Measurements conducted using dielectric resonator at 10 GHz, which is a critical tool for optimization of coating conditions.
MgB2 films coated on mockup 1.3 GHz Cu RF cavities by HPCVD with as high as 37.6 K measured by AC susceptibility. Ready to coat a real Cu cavity.
Acknowledgment
Ann-Marie Valente-Feliciano, Thomas Jefferson National Accelerator Facility
Steven Anlage, University of Maryland
Funding from US Department of Energy, Office of Science

## Slide 17
, ,
Estimation of RF Rs from residual resistance