## Slide 1: PoPLaR Phase 2: RCF Analysis
Calvin Dyson

## Slide 2: Overview
RCF was used to estimate the dose delivered to the cells
Two Main Sources of Uncertainty
    Temporal (Variation in dose between shots)
    Spatial (Variation in dose across the cell dish)
Cu Scatterer (10𝜇m)
Mylar (50𝜇m)

## Slide 3: Dataset
Collected 50 Films
    20 larger than the cell dish
    30 smaller than the cell dish
Used the larger films to estimate the conversion of the smaller films to the full cell dish area
Range of shots taken: 1-14
    Key shot numbers are 4, 7 and 14 shots
Range of Dose Delivered: 0.01 – 9.99 Gy
Notes: Mention:
Large Variation No Clear dependence on shot number
Potential Dependence on Smaller Film Size

## Slide 4: Calibration
Models
  Gafchromic’s Recommended Model
  Polynomial Model

## Slide 5: Calibration
Evaluation
Model | a | b | c | 𝜒r2
Gafchromic | 2.13±0.13 | 0.167±0.015 | 2.12±0.13 | 9.87
Polynomial | 6.43±0.33 | 43.7±8.8 | 2.85±0.25 | 9.63
Polynomial (C fixed) | 6.36±0.23 | 41.3±2.4 | 2.78 | 8.55
Devic S et al 2004 Dosimetric properties of improved GafChromic films for seven different digitizers Med.
Phys. 31 2392–401

## Slide 6: Estimating Dose Delivered
Performed a Maximum Likelihood Fit on the mean dose and the shot-to-shot variation
Assumes normally distributed shot-to-shot variation.
Ni = Number of Shots, 𝜇 = Mean Dose per Shot,
𝜎shot = Shot-to-shot Variation, 𝜎D = Measurement Error
Result:
    𝜇 = 0.57 Gy
    𝜎shot = 0.40 Gy
    𝜎D = 0.03 Gy
    𝜒red2 = 1.05 (p=0.384)
Shot-to-shot Variation

## Slide 7: Estimating Dose Delivered
20 Large Films
Mean CV: 13.93%
No obvious dependence on mean dose
Spatial Variation

## Slide 8: Estimating Dose Delivered
Compare Dose Variation Sources
Main Error is shot-to-shot variation
Need to reduce the spatial variation as well
CV becomes the bigger proportional error at 26 shots (or 14.8 Gy)

## Slide 9: Estimating Dose Delivered
Final Results
Shots | Predicted Dose /Gy | Uncertainty /Gy
4 | 2.28 | 0.87
7 | 3.98 | 1.21
14 | 7.97 | 1.87

## Slide 10: Uniformity Improvement
Move Copper Scatterer Closer

## Slide 11: Uniformity Improvement
Move Copper Scatterer Closer

## Slide 12: Conclusions and Future
Shot-to-Shot Variation is the biggest uncertainty
    Requires in-beam diagnostic
Delivered roughly twice the dose as expected
  Spatial Non-Uniformity Considerable
      Moving the scatterer closer does improve uniformity
      Delivers a little less than half the dose per shot
12