## Slide 1: Target Systems Vacuum Vessel and Feedthrough Design ISIS/CERN collaboration progress workshop 1rst - 2nd June 2026
Gemma HUMPHREYS - Jean-Louis GRENARD – CERN – SY-STI-TSP
On Behalf of the HI-ECN3 Project Team

## Slide 2: Contents
Facility Overview
Contents and functions
Objective and requirements for the study
Design Development
Installation and Alignment
Failures
Vacuum system
Feedthroughs
Manufacturing
Vacuum Vessel Seal
Prototype Plans
Opening Sequence

## Slide 3: Target Station & Target system
P+
Target System
Copper Proximity Shielding
Trolley assembly
Vacuum vessel
Magnetised Hadron Stopper
Target bunker (iron & concrete)
SHiP
Experiment
9m
5m

## Slide 4: Contents and Functions
Primary Functions:
Minimise activation of the surrounding air and corrosion risks through maintaining vacuum level of 10-3 mbar
House the target, related proximity shielding and relevant cooling components
Maintain target alignment with the beamline
Auxiliary Functions:
Contain a system to ease the removal of the target and shielding
Provide structural integrity
Integrate with the safety systems
Allow for integration of coolant pipes into the vessel
Retention vessel if coolant leakage occurs
Internal Vacuum Vessel Systems Cross Section
Internal load: ~150 tonnes
Empty Weight: ~13 tonnes
Prompt Dose Rate: ~1.6E11 uSv/h
Internal Volume: ~30m3

## Slide 5: Objective of Preliminary Design Study - Streicher
16-week mechanical design contract for the preliminary design of the vacuum vessel system
  Expertise in the field of competence (design, manufacturing / construction, installation)
Further develop the CERN concept with design and structural calculations
Develop the vacuum system
Understand risks associated with vacuum vessel
Develop the opening sequence and any tooling required
Provide standards and regulations to be followed
Provide qualification criteria for welds and vacuum acceptance
Generate an estimate of the cost for the final production of the vacuum vessel
Price enquiry and specification: https://edms.cern.ch/document/3315586/1

## Slide 6: Requirements
Fit in a defined envelope within target station shielding
Comply with material restrictions
  Low Cobalt content
  No magnetic properties
Fit all required components within the vessel
Locate all possible failure points on one part of the vessel
Handle if a water leak was to occur in the vessel
  Drainage options
  How to dry vessel

## Slide 7: Design Change from Previous Year
6579mm
2050mm
2941mm
TSAC1
Internal volume: 28.96m³
Weight: 25 Tonnes
6555mm
2060mm
2950mm
TSAC2
Internal volume: 26.52m³
Weight: 13 Tonnes
Full Structural assessment
Method of installation and alignment
Door installation and opening
Developed shielded lifting tool
Vacuum pumping system
Drainage connections
Guide Rails

## Slide 8: Shielded Lifting Tool
5cm thick carbon steel tool
Moved into place to provide some shielding when door is being opened
Utilises manual crank to move towards the door
  Design adaptation to include robotic handling
Locking pins secure tool to door
Lead blankets for specific operations
Shielded Lifting Tool
Cut-outs for feedthroughs
Manual crank
Connection to vessel door
Rollers
Lead Blankets

## Slide 9: Water Leak Scenario
Notes
Chamber State
Action
Step
Leak detection system in
cooling system
Under Vacuum
Leak detection (pressure rise
-
> chiller
stops)
1
To prevent larger leaks and allow
evacuation
Atmosphere/under vacuum
Dry the cooling loop (collect water to
chiller, seal inlet/outlet)
2
Only done for large leaks (>few
litres
)
Under Vacuum
Pump drainage (water pump via
drainage network)
3
Usage of vacuum pumping
system
Under Vacuum
Vacuum drying residual water left in tank
4
Water Collection
Drainage Pipe
Gutter System

## Slide 10: Installation and Alignment
Drilling and installing anchor bolts into TCC8
Base plates installed and vertical alignment performed with set screws and laser trackers (0.2 - 0.3mm accuracy)
Base plates grouted with high-compressive strength cement-based compound
Vacuum chamber installed onto base plates
Chamber positioned with chamber alignment bolts (±20mm)
Chamber fixed to base plates
Size of base plates ~1m2 to remain within the 30 tonnes/m2 average overall weight limit in TCC8
Ground Anchor Bolts
Chamber Alignment Bolts
Levelling Bolts
Levelling Bolts
Laser Tracker Interface
Openings for Concrete Pouring
Base Plate Design

## Slide 11: Failure Reduction and Recovery
Cooling pipes have welded connections to flanges
  Reduced chances of leakage
All components mounted onto the vacuum vessel door
  One replaceable component if there were to be failures
Weld inspection
  All welds on the pipes check with x-ray and die penetrant testing
  Pressure tests performed
Pressure relief valve + burst disk – used in case of a helium leak from target cooling
  Redundancy on overpressure scenarios
Event of failure inside of vessel:
  Send a robot inside
  Mount a shielded cabin onto the rails and put inside to carry out maintenance

## Slide 12: Failure Impacts to Surroundings
Leak in the water or helium cooling within vacuum vessel
Large vacuum leak in vessel
With proton beam window
Without proton beam window
Contamination spreads out of vacuum vessel into upstream beamline
Loss of upstream beamline vacuum
Contamination remains within the vacuum vessel
Upstream beamline remains under vacuum
Could foresee a fast vacuum valve and/ or proton beam window

## Slide 13: Vacuum System
Rotary Oil Pump + Roots Booster |  | Dry (screw) Pump + Roots Booster |  | 2 Dry Multi-Stage Roots Pumps + Roots Booster | 
Ultimate pressure lower than dry pump | Possible oil pollution in primary pump with radioactive dust – preventable with soft start pumping | No possible contamination of chamber with oil | Worse vacuum performance | No possible contamination of chamber with oil | Worse vacuum performance
No contamination of chamber from oil vapors |  |  | More complex construction and servicing | No annual servicing | More complex construction and servicing
Annual oil change on site |  | Annual oil change, belt replacement, full cleaning on site |  | Major service at supplier premises every 4 years | 
Major service at supplier premises every 10 years |  | Major service at supplier premises every 4 years |  |  | 
Final selection still to be fully validated by BE-EA

## Slide 14: Vacuum Vessel Preliminary P&ID
Outlets both have filters and are connected to the ventilation system
Filters and outlets could be grouped together
Water drainage connected to contaminated water network
Proposed systems to be checked for compatibility with CERN NA primary vacuum infrastructure

## Slide 15: Vacuum Vessel Feedthroughs
Action description
Action
needed
before
door
removal
Additional info
Component description
Type of gasket
Placement
Connection
type
Function
No.
Overpressure relief valve should be outside radiation zone.
There is O
-
ring in the body of the valve
–
degradation would
cause lose of sealing. No piping connected on the outlet
–
only particle filter.
NO
Relief pressure =
±
0,15
bar(g). Flow rate max =
210 l/min. Added particle
filter on the outlet
FKM/EPDM
Vacuum piping
DN40 ISO
-
KF
Pressure relief
valve
1
Can be connected on the door or on the main flange, no
maintenance needed (metallic sealing). No piping
connected.
NO
Bursting pressure = 0,45
–
0,48 bar(g). Flow rate >
500 l/s. Free volume of
the chamber = 7500 l
Metallic
Door/Flange
DN63 ISO
-
K
/ ISO
-
CF
Pressure burst
disc
2
Vacuum system can be connected to the main flange
–
then
there would be no action needed. If connected to the door
–
then piping
must
be dismounted before door opening.
YES/NO
10
-
3
mbar to be
reached
EPDM if it is
located on the
door. Metallic if it
is located on the
flange
Door/Flange
DN100
Vacuum system
3
No action needed, outside of the radiation zone.
NO
Manual valve with LOTO
function + Particle Filter
FKM/EPDM
Vacuum piping
DN40 ISO
-
KF
Venting
4
Disconnection of communication cables.
YES
2
×
Pirani measuring tube.
Measuring range: 8E
-
4
–
1000 mbar
Metallic
Door
DN16 ISO
-
KF / ISO
-
CF
Vacuum gauge
5
To be determined
Door/Flange
Level limit switch
/ humidity sensor
6
No action needed.
NO
Water collection channel
connected to a swagelok
fitting on the main flange
Metallic
Flange
3/4"
Swagelok
tube fitting
Drainage system
7
Unplug thermocouples from door
Yes
SUB
-
D (2 x 50 Pins)
Metallic
Door
DN100
ISO
-
K/ISO
-
CF
Thermocouples
8

## Slide 16: Incorporation of Cooling Feedthroughs
Target Helium Cooling
Proximity Shielding water cooling
Extra feedthroughs for instrumentation
Bellows connected to vacuum vessel door
Assessing to include remote flange connection
Design study ran with Destec to incorporate remote flange connections onto the vessel door
Multiple configurations assessed
Sealing both the pipes and the vacuum proved challenging with added weight of heavy connectors
Complexity outweighed the benefits
Integrated dose: 1-5 kGy/year in this area
Residual dose: ~500-800µSv/h compatible with human intervention

## Slide 17: Vacuum Vessel Cooling Feedthroughs
2026 design under development
2x Helium cooling pipes DN80
Thermocouples
2 x Water cooling pipes DN25
Support structure for pipes
Remote chain clamp
Water Cooling
Helium Cooling
Design courtesy of C.Mucher

## Slide 18: Manufacturing
Base of vacuum vessel is produced and reinforcements welded
Interface rails machined precisely and welded to the base
  Previous experience with similar sized vessel achieved ±0.2mm accuracy
Rest of vacuum vessel welded around the base
Sand blasting finish
Weld inspection
  Penetrant testing
  Ultrasonic testing
Leak test
  Performed before shipped and upon installation
Interface Rails
See L.Jorat talk for more information on vacuum vessel rails
Lip to ensure alignment when installing trolley rails

## Slide 19: Vacuum Vessel Seal
Vessel door, helium cooling and water cooling all require elastomer sealing
Investigations into the radiation effect to seals ongoing
Seals to be replaced every time the vacuum vessel is opened – once every 5 years
Radiation to materials level in seal locations reaches up to 3kGy/year
Seals must withstand at least 15 kGy total

## Slide 20: Elastomers for Vacuum Vessel Seals – Overview
Motivation | Develop know-how in testing elastomers at CERN Assess radiation effects on elastomer materials
Materials | 3 silicone rubbers,1 EPDM Cylindrical samples and O-rings
Irradiation | 37 - 201 days of irradiation at nTOF NEAR (neutron-dominated field) 5 dose points between ~170 - 940 kGy 1-, 2-, and 3-month irradiations finished in 2025 5- and 7-month irradiations ongoing (until end of August 2026)
Pre- & post-irradiation tests | Leak tightness Residual deformation after 10% compression Surface aspect (microscopy)
Slides courtesy of D. Zemkova
Compression Assembly
O-ring Assembly

## Slide 21: Elastomers for Vacuum Vessel Seals – Project status
Materials & tools | Samples provided by an external company Compression assemblies and tools provided by an external company and CERN
Sample preparation | Pre-irradiation tests done at CERN and the external company Irradiation assemblies prepared at CERN
Tests | Tests on 2025 samples scheduled in May 2026 Tests on 2026 samples in Autumn/Winter 2026
Results | Preliminary results to be presented in June 2026 Final results to be published in collaboration with the external company
Slides courtesy of D. Zemkova
Helium Leak Testing

## Slide 22: Prototype Plans
Feedthrough Connections
  Mock-up to assess the bolting of the flange to the door and chain clamp operations
  Assess robotic compatibility with chain clamps
  Ensure pipes can be well managed during disconnection/connection process
  Check spacing of pipes is compatible with robotic intervention
Vacuum vessel seal irradiation testing
No prototyping for the vacuum vessel
  Product proven to work in industry already

## Slide 23: Vacuum Vessel Opening Sequence
21 mSv/h, 4 hours cooldown time
Courtesy of C. Mucher

## Slide 24: Beam Line Removal

## Slide 25: Beam Instrumentation Package Removal

## Slide 26: Collimator/Mask Shielding Removal

## Slide 27: Colllimator/Mask Removal

## Slide 28: Beam Pipe Removal

## Slide 29: Feedthrough Disconnection

## Slide 30: Chain Clamp Removal

## Slide 31: Pipe Removal
Contamination risk during disconnection

## Slide 32: Second Pipe Removed

## Slide 33: Pipe Support Removal

## Slide 34: Thermocouple Removal

## Slide 35: Flange Bolt Unscrew

## Slide 36: Water Pipes Disconnected and Removed

## Slide 37: Door Unbolted

## Slide 38: Lifting Tool Transported to Chamber Guide Rails
Courtesy of Streicher

## Slide 39: Manipulation of Lifting Tool to Chamber Door
Can also be performed by robotics

## Slide 40: Connection of Lifting Tool to Door

## Slide 41: Manipulation of Lifting Tool + Door to End Stop

## Slide 42: Transporting Lifting Tool + Door to Storage Frame

## Slide 43: Separating Door from Lifting Tool
Clamp door to storage frame
Lift levers to separate tool from door
Crank tool away from door

## Slide 44: Conclusions
Preliminary design of vacuum vessel completed (see additional documentation)
  Awaiting final report
Risk analysis ran in parallel by Streicher (see additional documentation)
  Fundamental risks of vacuum chamber
  Basic control requirements
  Identification of residual hazards
Second contract for the detailed final design and manufacture of the vacuum vessel and its associated system in 2027
Feedthrough design ongoing
  Prototype to be developed for the connection/disconnection