## Slide 1: Characterization of the Venturi flowmeters for the control of ITER magnets
S Varin1, J–M Poncet1, J André1, E Ercolani1, N Luchier1, C Mariette1, N Clayton2, J-Y Journeaux2
1 Univ. Grenoble Alpes, CEA, INAC-SBT, 38000 Grenoble, France
2 ITER Organization, Route de Vinon-sur-Verdon, 13115 Saint Paul-lez-Durance, France
presented by S. Varin

## Slide 2: Introduction
Context
June 2014: Supply contract between CEA-SBT and the Magnet Division of the ITER Organization
CEA/SBT was in charge of the design, manufacturing & delivery of 277 flowmeters
June 2018: Final delivery
Flowmeters will be used to operate the superconducting magnet system of the ITER tokamak

## Slide 3: Introduction
Design, manufacturing & delivery of 277 flowmeters
Reynolds number not included in the standard NF EN ISO 5167-4 (2×105 < Re < 1×106)
Experimental characterisation
 Manufacturing strategy
Size | Mass flowrate | Temperature | Pressure | Reynolds number | Number
DN8 | 0.1 to 1 g/s | Near 300 K | 3.8 to 4 bar | 6×102 to 6×103 | 18
DN10 | 1 to 7 g/s | Near 300 K | 3 to 4 bar | 3×103 to 3×104 | 42
DN15.20 | 2 to 20 g/s | 4.2 to 6 K | 4 to 10 bar | 4×104 to 5×105 | 89
DN15.30 | 3 to 30 g/s | 4.2 to 6 K | 4 to 10 bar | 6×104 to 8×105 | 28
DN20 | 13 to 130 g/s | 4.2 to 6 K | 4 to 10 bar | 2×105 to 3×106 | 40
DN25 | 40 to 400 g/s | 4.2 to 6 K | 4 to 10 bar | 4×105 to 6×106 | 60
Operation of the current leads at room temperature
Control of the supercritical helium flow in the magnets

## Slide 4: Sizing and manufacturing strategy
Sizing
21°
D
d
d
7°
Flow
  Where:
  : mass flowrate, density of the fluid, pressure at the neck diameter
  upstream and neck diameters
  : pressure and temperature in the upstream pipe
  : flow coefficient, with the expansion factor and the discharge coefficient
Designed following the standard NF EN ISO 5167-4

## Slide 5: Sizing and manufacturing strategy
Measurement accuracy
  Where:
  The subscript m refers to the uncertainties resulting from the measurement at CEA/SBT
  : mass flowrate
  upstream density
  upstream and neck diameters
  : upstream pressure and temperature
  : differential pressure between the upstream section of the Venturi tube and its restriction

## Slide 6: Sizing and manufacturing strategy
Manufacturing strategy
277 flowmeters  avoid an individual characterization
Tight dimensional tolerances
Reproducible behaviour
Sampling measurement

## Slide 7: Sizing and manufacturing strategy

## Slide 8: Method for the characterization of flowmeters
Warm bench
Determination of the flow coefficient by comparison with a Coriolis flowmeter

## Slide 9: Method for the characterization of flowmeters
Cold bench
Loop cooled by a helium refrigerator producing 800 W @4.5 K
Test of 9 Venturi tubes
Determination of the flow coefficient by comparison with a Coriolis flowmeter

## Slide 10: Results
Warm flow test
Test of 7 DN8 and 10 DN10 flowmeters
Influence of the pressure and differential pressure (mass flowrate) on the flow coefficient
Difference between flow coefficients of a given type of flowmeters < 2% at full-range

## Slide 11: Results
Cold flow test
Test of 9 flowmeters
No influence of the pressure, temperature and mass flowrate
Difference between flow coefficients of a given type of flowmeters < 3% at full-range

## Slide 12: Validation of the manufacturing strategy
Achievement of the objectives
Determine the flow coefficient
Constant values for cold flowmeters
Flow coefficient fits as a function of the differential pressure for warm flowmeters
Validate the manufacturing strategy
Reproducible behaviour (difference between flowmeters of a given type <3% at full-range)

## Slide 13: Validation of the manufacturing strategy
Probable error on the measurement of the mass flowrate
Influence of the mass flowrate on the measurement accuracy
5.2
2

## Slide 14: Conclusion
Conclusion
Test of 17 warm flowmeters & 9 cold flowmeters
Determination of the flow coefficient
Constant for cold flowmeters
Flow coefficient fits as a function of the differential pressure for warm flowmeters
Reproducible behaviour
 Validation of the manufacturing strategy
Good accuracy
 2% at full-range
(except in the neighbourhood of the critical point)
The technical specifications were met for all flowmeters