## Slide 1
Superconducting Magnets
Mike Courthold
Cryogenics and Magnetics Group
Applied Science Division

## Slide 2
SC coils consume virtually no power, although they consume refrigeration power
  energy costs are lower (~10x for large magnets)
SC coils achieve high current densities
  compact windings
    windings less obtrusive
    more complex designs
      easier to achieve very uniform fields
  higher fields, without the need for iron (use more turns)
    reduced bend radius for accelerators
      smaller accelerator rings
      reduced capital costs
      new technical possibilities
  higher field gradients
    higher quadrupole gradients
All the above make physicists very happy
Why use Superconducting Magnets?

## Slide 3
The highest engineering current densities are not easy to achieve (see later)
SC coils invariably take much longer to design, manufacture, and implement (7 years for Delphi)
  Start-up costs and effort are significantly greater.
  Need to be used for longer to achieve overall cost-savings (9 years for Delphi)
SC coils must be cooled, typically to LHe temps by a refrigerator or cold heads
  Extra complexity and cost
  SC coils must be firmly supported to resist the strong magnetic forces, as well as gravity, but thermally isolated from their surroundings
    Important to get this right
      Tiny movements can induce a Quench
      Heat leaks can prevent SC operation or reduce the margin
  2-stage cooling needed
    1st stage – Cool radiation shields; and intercept and extract heat from current leads, instrumentation wiring, etc., which would otherwise undermine the 2nd stage cooling
    2nd stage - Cool and maintain cold-mass at typically 4.2 K
What’s the catch? 1

## Slide 4
Example Magnets – Delphi Solenoid
Delphi
1.2 T
5.5m dia
6.8m long
110MJ

## Slide 5
Example Magnets - Atlas

## Slide 6
Example Magnets – Atlas End Cap Toroids

## Slide 7
Typical Refrigerator / Cryo-Cooler Options

## Slide 8
Quench detection and action must be:
  Comprehensively designed and analysed
  100% reliable, and over-engineered, but not over-sensitive – challenging!
SC magnets can easily be damaged by careless or inadequate care, monitoring, maintenance.
  Faults and issues with SC coils can be very difficult and costly to rectify.
    9 mnths re MICE Decay Solenoid current-lead failure during Quench Detector upgrade
    9 mnths re Delphi Quench detection system, due to design fault creating coil movement when ramping
    Several mnths re Delphi current-lead insulator failure, due to excessive cold shock
    9 mnths re Aleph cryostat repair, following uncontrolled Quench, due to UPS failure
    1 yr re LHC beamline repair, following uncontrolled Quench of multiple SC beamline magnets
What’s the catch? 2

## Slide 9
This is a feature of all Low Temperature Superconductors (LTS) (see picture)
It is the boundary between superconductivity and normal resistivity in 3-dimensional space
Everywhere above/below the surface is resistive/superconducting respectively
Each alloy exhibits three important characteristics
  Bc – the critical field at zero temperature and current
  Tc – the critical temperature at zero field and current
  Jc – the critical current density, which depends on how it has been processed
LTS Properties – The Critical Operating Surface 1

## Slide 10
LTS Properties – The Critical Operating Surface 2

## Slide 11
LTS cable is usually operated at 4.2 K, so the Critical Surface is often reduced to a more-convenient curve of critical current (density) versus field at 4.2 K (see curve)
Low Temperature Superconductors (LTS)
  NbTi
    The alloy of choice for over 60 years
    It is a ductile alloy (the only one)
    Critical Current Bc is 9 T
  Nb3Sn
    Much higher critical current versus field than NbTi at 4.2 K
    Critical Current Bc is 16 T
    Unfortunately it is an intermetallic compound, and thus brittle, with poor mechanical properties
LTS Properties – The Critical Line at 4.2 K
Martin Wilson

## Slide 12
  LTS wire comprises very fine SC filaments embedded in high purity copper stabiliser
    To reduce flux jumps (magnetic instabilities)
    To minimise field distortions due to SC magnetisation
    NbTi wires are typically 0.3 – 2 mm diameter, comprising 30 – 56 filaments, each 30 - 150 um diameter.
    Nb3Sn wires are typically 0.5 - 1 mm diameter, comprising 24600 filaments, each 1.5 – 3.3 um diameter.
    The wires are twisted together like a rope, to form a cable (“Rutherford Cable”)
      To reduce inter-filament coupling and AC losses
    The cable is embedded in high purity aluminium
      To protect the SC following a Quench
      To reduce flux jumps
LTS Properties – Cable Composition 1

## Slide 13
LTS Properties – Cable Composition 2

## Slide 14
LTS Properties – Typical use of “Rutherford Cable”

## Slide 15
There are many High Temperature Superconductors (HTS) with Critical Temperatures conveniently above 90K, eg YBCO (YBa2Cu3O7)
  Will thus operate in liquid nitrogen
    Considerably cheaper and less complicated than LHe
HTS features a disappointing “Flux Flow” operating region, in which the material is resistive, whilst still superconducting, rather than a well defined Critical Temperature
  HTS can still be used in regions of both high temperature and high field, but current density is a problem
HTS is produced as a tape
  Frequently used for current leads within cryostats that use cold-heads
    But temperature and field must be well monitored & controlled
      For magnets utilising both HTS and LTS cables, HTS cables MUST have adequate Quench protection, including substantial current shunts
    Minimises thermal load on the cold-head 1st stage
      Increases efficiency of cold-head 2nd stage
    Not needed with refrigerators, which cool current leads with return He gas
HTS Properties

## Slide 16
  When designing a magnet, the Engineering Current Density (Jeng) is the most important characteristic
    The resultant field and hence the actual current is all that really matters
  Engineering Current Density takes account of:
    The actual proportion of SC in each wire (fill factor), typically 0.4 to 0.2
    The fill factor of the wires in the coil pack, typically 0.7 to 0.8
      (including insulation, cooling channels, stabiliser, mechanical reinforcement, etc)
    Thus Jeng is realistically only 15% - 30% of Jsc
      The actual coil pack volume is inversely proportional to Jeng/Jsc
Engineering Current Density

## Slide 17
  A magnet Quenches when it irreversibly changes from the superconducting to resistive state
    Consequently, all the stored energy is converted into heat
      through resistive heating of the SC cable, which is dissipated in the cold mass
      with the option of extracting energy via external dump resistors
        NB: The power circuit must be designed with care, if HTS current leads are used
      All the liquid cryogens will also be rapidly expelled via relief valves
  Superconductors have a very high normal-state resistivity
    If Quenched, a very high temperature would be reached very quickly
      The coil might suffer damage and/or delaminate at temperatures above 150K, and certainly above 250K
    By embedding the SC filaments in a matrix of copper
      The matrix helps to stabilise the SC against flux jumps
      The matrix allows current to redistribute in the low-resistivity matrix
        The peak temperature will be lower
        The Quench detection system has more time to react and act on the power circuit
Quenching

## Slide 18
Magnets should be designed to Quench very rarely, as each Quench causes stress
The anticipated Quench behaviour of the magnet must be designed and analysed meticulously
Quench protection circuits comprise some or all of the following:
  Voltage taps that measure and compare the voltages between comparable parts of the coil circuit by means of sophisticated Wheatstone Bridges
    These typically trip the power-circuit when a persistent imbalance of 1V is detected
    The detection system must be 100% reliable, but avoid false signals – challenging!
      Use hard-wiring, PLC, or Linux (in order of merit), but never Windows
      NB: All such trips will result in an actual Quench, whether or not the signal was real
    The detection system must compare overlapping segments, to avoid excluding joints
  Pressure, temperature, and vacuum sensors, to react to Quench precursors with a “slow trip” and avoid an actual Quench where possible
Whenever a Quench is initiated/detected, it is imperative to force all circuits to Quench together, by means of Quench-heaters, to avoid hot-spots and/or unbalanced forces
Quench Protection

## Slide 19
  Rapid current-ramping
    Causes eddy-current heating => Quench
  Exceeding the Critical Current => Quench
    Avoid exceeding 80% of the maximum current (Short-Sample Current)
  Coil or current lead movements within the magnetic field
    Can produce an artificial Quench signal
  Cracking of the resin or flux jumping
    Causes a transition, which allows some current to flow in the stabiliser
      Can lead to a Quench, if not stabilised quickly enough
  Resistive joint, creating heat
  Thermal leak of the cryo-system
  Training
    All SC magnets undergo some “training”, whereby they progressively approach their theoretical Critical Current
      Better-designed magnets suffer less training
      NB: Warming-up a magnet can undermine previous training
Quench Causes

## Slide 20
  Static Forces
    Suspension and tie rods are required to handle gravitational loads, as well as magnetic forces
      The rods have to be strong enough to handle the forces
      The rods must also be thermally insulting between the cold mass and its support frame
  Forces during a Quench
    Significant unbalanced forces can develop during a Quench due to:
      Unbalanced or partly-delayed collapsing fields
      Forces between the Quenching magnet and any neighbouring magnets
Forces within Magnets

## Slide 21
  Fringe fields produced by a non-shielded magnet can adversely affect:
    Neighbouring instruments and equipment
    The magnet’s own instrumentation, including its protection system
    People in the vicinity
    NB: The presence of ferrous objects can strongly enhance the local field beyond what might be predicted in air
Beware Fringe Fields

## Slide 22
Thank you
technologysi.stfc.ac.uk
@STFC_Matters
Science and Technology Facilities Council
Science and Technology Facilities Council