## Slide 1: EPAC operations: Diagnostics, alignment and shot mode

## Slide 2: Content
Operational laser modes
Daily start up and alignment procedure
Full power shot checks
Repetition rate change procedure
Diagnostics
Auto-alignment

## Slide 3: Operational laser modes
Experience gained from Gemini operations has shown that the users require access to range of energies to setup, characterise and perform their experiments
To provide a range of energy to the users, the laser will be operated under a selection of power modes
4 modes have been defined that range from μJ to the J level
Each mode offers a selection of activities that can be undertaken in the Experimental Areas

## Slide 4: Operational Laser Power Modes
Mode 3
1% of Front End energy to allow limited procedures in experimental areas
Mode 1
Maximum energy of system
Mode 2
Maximum energy from Front End
Mode 4
0.05% of Front End energy to a level that it is ‘safe’ to work in experimental areas
Potential for creating radiological hazard
Laser hazard
Used for data taking
Used for testing/calibration
Used for tuning and timing
Used for alignment

## Slide 5: EPAC daily start-up and alignment
Alignment for each section follows a standardised procedure:
  Verify the alignment of the section with its internal CW
  Receive pulsed beam from previous section
  Match the pulsed beam to input references in the current section
  Verify pulsed beam alignment through the section
Each section can be aligned in parallel
  The alignment is done with CW first
  Each section of the system has CW lasers and fixed input references
    Reliable and stable input beam
Matching and handover
  Beam from previous area is matched to the fixed input references using pairs of dedicated mirrors

## Slide 6
EPAC daily start-up and alignment
Front end seed and ps OPA
TI:S CW
verification
Modes 2 3 4
available
(low power)
Compressor verification with
FE beam
Ti:S verification
with FE beam
Align FE to
Ti:S input diagnostics
Ti:S full power checks
Compression measurement
Pump FE
verification
Compressor
CW verification
10J verification
EA internal
CW verification
Pump laser IR ready
100J verification
Wavefront
measurement
Full power
run-up
Full power through compressor
Compressed
full power
checks
Warm up front end pumps
Warm up front end oscillator
Start Pump
cryostats
Remote warm-up
Front end
ns OPA
Verify pulsed beam to EA input diagnostics
SHG verification
Low power verification
Full power available for shot start procedure
Output diagnostics verification
Compression and wavefront measurements
Front-end
Ti:Sapphire
Verify CW to EA input diagnostics
CW available
CW verification
Pump
Experimental area

## Slide 7
Front-end
Ti:Sapphire
Pump
Experimental area
Front end checks
Output energy correct and stable
Output beam profile good
Output spectrum shape good/stable
Output pointing good/locked
Ti:S multi-pass pulsed checks
Align pulsed beam through multi-pass
Check wavefront and optimise
Output beam energy stable
Output beam profile good
Pinhole beam centricity good
Lock output pointing
Ti:S multi-pass initial checks
Switch on Ti:S crystal cooling
Check CW alignment of Ti:S multipass
Set VSF lenses for HP mode
Set AO for HP mode
Short pulse checks
Check compressor CW alignment
Check compression
PFT checks
Optimise post comp wavefront
Pump FE verification
Front end/pre amp start up
Check front end alignment
Front end pulsed mode/energy ramp
10 J verification
Switch on/Stabilise 10 J pumps
Check 10J alignment - CW only
Check overlap 10J CW and FE pulsed
100J verification
Search pump room
Switch on/ Stabilise 100J pumps
Check 100J alignment - CW only
Check overlap 100J CW and TJ pulsed
Set CW pump beam for SHG alignment
Check CW IR pump beam to LBO crystal
Check alignment of pulse beam at ~5 J
Open diverter
Align to the Ti:S 20% efficiency ~1 J green
Close diverter to green dump
Ramp up pump for required green energy
Output energy correct/ Check pulse shape
Set multi-pass VSF lenses for de-timed alignment
Pump Ti:S with beams de-timed
Allow fully pumped Ti:S crystal to hit thermal stability
mode
Open wall shutter to pump room
SHG checks
Align beam using Green CW to Ti:S
Turn on Dark field through SHG
Check/set Ti:S cooling for FP
Close wall shutter from front end
Move in Ti:S output beam diverter
Ti:S full power checks
Check HP alignment of multi-pass
Set diagnostics for FP
Set VSF lenses for timed FP
Time pump beams for FP
Check FP beam energy to be correct
Measure/set wavefront to be correct
Check beam NF profiles
Switch on output pointing
Compressed full power checks
Remove fast output beam dump
Propagate full Energy through compressor
Allow for thermal equilibrium
Set/measure pulse duration
Check PFT
Check/correct wavefront
Insert fast output beam dump
Insert post compressor beam dump
Remove fast output beam dump
Hand control to EA1
Open wall shutter to Target Area
Insert fast output beam dump
Remove post compressor beam dump
Remove fast output beam dump
Full power shots
Alignment checks
NF & FF into area are good (FP leakage diagnostics)
Activate beam pointing stabilisation
Activate defocus stabilisation
Turn on gas
(Start tape drives)
EA1
Post shot wind down procedures
EPAC full power shot checks

## Slide 8
Front-end
Ti:Sapphire
Pump
Experimental area
Front end checks
Output energy correct and stable
Output beam profile good
Output spectrum shape good/stable
Output pointing good/locked
Ti:S multi-pass pulsed checks
Align pulsed beam through multi-pass
Check wavefront and optimise
Output beam energy stable
Output beam profile good
Pinhole beam centricity good
Lock output pointing
Ti:S multi-pass initial checks
Switch on Ti:S crystal cooling
Check CW alignment of Ti:S multipass
Set VSF lenses for HP mode
Set AO for HP mode
Short pulse checks
Check compressor CW alignment
Check compression
PFT checks
Optimise post comp wavefront
Pump FE verification
Front end/pre amp start up
Check front end alignment
Front end pulsed mode/energy ramp
10 J verification
Switch on/Stabilise 10 J pumps
Check 10J alignment - CW only
Check overlap 10J CW and FE pulsed
100J verification
Search pump room
Switch on/ Stabilise 100J pumps
Check 100J alignment - CW only
Check overlap 100J CW and TJ pulsed
Set CW pump beam for SHG alignment
Check CW IR pump beam to LBO crystal
Check alignment of pulse beam at ~5 J
Open diverter
Align to the Ti:S 20% efficiency ~1 J green
Close diverter to green dump
Ramp up pump for required green energy
Output energy correct/ Check pulse shape
Set multi-pass VSF lenses for de-timed alignment
Pump Ti:S with beams de-timed
Allow fully pumped Ti:S crystal to hit thermal stability
mode
Open wall shutter to pump room
SHG checks
Align beam using Green CW to Ti:S
Turn on Dark field through SHG
Check/set Ti:S cooling for FP
Close wall shutter from front end
Move in Ti:S output beam diverter
Ti:S full power checks
Check HP alignment of multi-pass
Set diagnostics for FP
Set VSF lenses for timed FP
Time pump beams for FP
Check FP beam energy to be correct
Measure/set wavefront to be correct
Check beam NF profiles
Switch on output pointing
Compressed full power checks
Remove fast output beam dump
Propagate full Energy through compressor
Allow for thermal equilibrium
Set/measure pulse duration
Check PFT
Check/correct wavefront
Insert fast output beam dump
Insert post compressor beam dump
Remove fast output beam dump
Hand control to EA1
Open wall shutter to Target Area
Insert fast output beam dump
Remove post compressor beam dump
Remove fast output beam dump
Full power shots
Alignment checks
NF & FF into area are good (FP leakage diagnostics)
Activate beam pointing stabilisation
Activate defocus stabilisation
Turn on gas
(Start tape drives)
EA1
Post shot wind down procedures
EPAC full power shot checks

## Slide 9
Front-end
Ti:Sapphire
Pump
Experimental area
Front end checks
Output energy correct and stable
Output beam profile good
Output spectrum shape good/stable
Output pointing good/locked
Ti:S multi-pass pulsed checks
Align pulsed beam through multi-pass
Check wavefront and optimise
Output beam energy stable
Output beam profile good
Pinhole beam centricity good
Lock output pointing
Ti:S multi-pass initial checks
Switch on Ti:S crystal cooling
Check CW alignment of Ti:S multipass
Set VSF lenses for HP mode
Set AO for HP mode
Short pulse checks
Check compressor CW alignment
Check compression
PFT checks
Optimise post comp wavefront
Pump FE verification
Front end/pre amp start up
Check front end alignment
Front end pulsed mode/energy ramp
10 J verification
Switch on/Stabilise 10 J pumps
Check 10J alignment - CW only
Check overlap 10J CW and FE pulsed
100J verification
Search pump room
Switch on/ Stabilise 100J pumps
Check 100J alignment - CW only
Check overlap 100J CW and TJ pulsed
Set CW pump beam for SHG alignment
Check CW IR pump beam to LBO crystal
Check alignment of pulse beam at ~5 J
Open diverter
Align to the Ti:S 20% efficiency ~1 J green
Close diverter to green dump
Ramp up pump for required green energy
Output energy correct/ Check pulse shape
Set multi-pass VSF lenses for de-timed alignment
Pump Ti:S with beams de-timed
Allow fully pumped Ti:S crystal to hit thermal stability
mode
Open wall shutter to pump room
SHG checks
Align beam using Green CW to Ti:S
Turn on Dark field through SHG
Check/set Ti:S cooling for FP
Close wall shutter from front end
Move in Ti:S output beam diverter
Ti:S full power checks
Check HP alignment of multi-pass
Set diagnostics for FP
Set VSF lenses for timed FP
Time pump beams for FP
Check FP beam energy to be correct
Measure/set wavefront to be correct
Check beam NF profiles
Switch on output pointing
Compressed full power checks
Remove fast output beam dump
Propagate full Energy through compressor
Allow for thermal equilibrium
Set/measure pulse duration
Check PFT
Check/correct wavefront
Insert fast output beam dump
Insert post compressor beam dump
Remove fast output beam dump
Hand control to EA1
Open wall shutter to Target Area
Insert fast output beam dump
Remove post compressor beam dump
Remove fast output beam dump
Full power shots
Alignment checks
NF & FF into area are good (FP leakage diagnostics)
Activate beam pointing stabilisation
Activate defocus stabilisation
Turn on gas
(Start tape drives)
EA1
Post shot wind down procedures
EPAC full power shot checks

## Slide 10: Repetition rate change
The system will feature set points for various parts, to automate part of the process of changing repetition rate
After the set point changes, alignment is confirmed before full energy ramp-up
Can be done in parallel across various sections

## Slide 11
Repetition rate change procedure
Change pump transport lens position
Change LBO angle and temperature
Change pump timings
Change lens positions
Check cryostat cooling flow
Change AO settings
Change Pockels cell trigger rate
Change fast diverter operation
Front-end
Ti:Sapphire multipass
Pump
SHG
Full alignment confirmation
Full energy
run-up
Send pump to LBO and let green output stabilise
Change Ti:S cooling parameters
Adjust PFT compensator
Move Ti:S telescope lenses to set positions
Change astigmatism corrector settings
Change AO settings
CW and pulsed alignment check
Full energy run-up
Change pre-LBO waveplate setting
No change to OPA operation

## Slide 12: EPAC Diagnostics
The diagnostics perform range functions in EPAC
Depending on the diagnostic type it serves one or more of these functions
The Control System has overarching view of diagnostics to ensure each section of laser is operating to the specified requirements and safely

## Slide 13: Diagnostic role classification
Verification and performance
Setup and alignment
Machine safety
Automation
Beam specification
Known failure modes
Useful / Helpful
Pointing Camera
Dark-fields
FROG
Increased system stability

## Slide 14
Front End
TiSa Amp
Compressor
Experimental Area
Spatial
Spectral
Energy
Temporal
Control System
Spatial
Energy
Temporal
Spatial
Energy
Verification and performance
Setup and alignment
Machine safety
Automation
Data acquisition

t
E
Pump
Spatial
Spectral
Energy
Temporal
Spatial
Spectral
Energy
Temporal
System Performance

## Slide 15: Auto-alignment and fast stabilisation
EPAC features two methods to stabilise beam pointing
  Auto-alignment
    Uses cameras and motorised mirrors
    Uses the main beam
    Compensates slow drifts and variations
  Fast stabilisation
    Uses position-sensitive detectors and a piezo-driven mirror
    Uses 2 dedicated CW beams
    Suppresses vibrations and jitter fast than 10 Hz

## Slide 16: Auto-alignment strategy
Objectives
  Maintain pointing through the system during an operations day
  Compensate slow drifts (sub-Hz operating frequency)
The initial approach: less is more
  Design a stable laser and add loops in key locations to the automated system
  Become more familiar with the system and adapt
Can be upsized or downsized as we gather more experience
  Many mirror mounts include actuators for accessibility reasons, can be changed from remote-operated to automatic

## Slide 17: Auto-alignment implementation

## Slide 18: Auto-alignment implementation
Initial positions for mirror actuators and cameras have been defined
Auto-align algorithm already demonstrated in testbed systems for various scenarios of perturbations and calibrations
Notes: Mention redundant physical references as well

## Slide 19
EPAC auto-alignment and stabilisation
EPAC pump
Front-end
Ti:Sapphire
multi-pass
Compressor
Experimental
area
Interaction chamber
Closed feedback
loop
Closed feedback
loop
Closed feedback
loop
Closed feedback
loop
Closed feedback
loop
Auto-alignment
control system
Beam position on cameras
Actuators on mirror mounts
Beam
transport
Closed feedback
loop
Drift compensation

## Slide 20: Fast stabilisation strategy
Objectives
  On-shot stabilisation
  Suppress jitter and vibrations faster than 10 Hz
Dedicated CW beams used to stabilise faster than the system’s repetition rate
  900 nm CW to measure jitter at the Ti:Sapph multi-pass
  400 nm CW to measure jitter at the compressor and beam transport
Piezo-driven mirror with a 1 kHz operating frequency after the front-end

## Slide 21
EPAC auto-alignment and stabilisation
EPAC pump
Front-end
Ti:Sapphire
multi-pass
Compressor
Experimental
area
Interaction chamber
Closed feedback
loop
Closed feedback
loop
Closed feedback
loop
Closed feedback
loop
Closed feedback
loop
Position-sensitive detector
900 nm
CW
1 kHz piezo
actuator
Position-sensitive detector
400 nm
CW
Fast stabilisation
control system
On-shot stabilisation
(up to 1 kHz frequency)
Auto-alignment
control system
Beam position on cameras
Actuators on mirror mounts
Beam
transport
Closed feedback
loop
Drift compensation

## Slide 22: Conclusion
We have defined the operational laser modes of EPAC
We have defined a set of facility procedures focusing standardisation and enabling parallel tasks
  Daily start up and alignments
  Full power shot checks
  Repetition rate change procedure
We have an integrated approach to the diagnosis of the laser
We have an auto-alignment and stabilisation system under development