## Slide 1: Ionacoustic Work Package
LhARA Collaboration Meeting
8th February 2023

## Slide 2: Online dose monitoring system
LhARA: protons / light-ions
Minimize damage on healthy cells: on-the-fly, non-invasive, range verification system
Real-time exact dose deposition profile
Bragg peak localization
Solution: ion-acoustic imaging

## Slide 3
As the beam penetrates the tissue, it is scattered & absorbed
Increase in temperature -Thermoelastic expansion
Increase in pressure - Acoustic wave emitted (thermoacoustic effect)
Detected by ultrasound detector
Image reconstruction
Ionacoustic imaging
Figure 1: Schematic of photo-acoustic imaging

## Slide 4
Simulation pipeline

## Slide 5
BDSIM

## Slide 6
source
(python file)
SmartPhantom
LION beamline - BDSIM
Figure 2: Render of LION beamline at Ludwig-Maximillian University Munich in BDSIM
2x quadropole magnets for focusing

## Slide 7
60.34 mm
55.77 mm
1844.7 mm
100 mm
332 T/m
318.5 T/m
10 mm bore diameter
LION beamline - BDSIM
Figure 3: Side-on view of LION beamline in BDSIM
2x quadropole magnets for focusing

## Slide 8
Generate angular distribution
source.txt
Source
Figure 4: Energy distribution of protons from source – simulated and analytic closely match

## Slide 9
Problems:
Simulating 1e7 particles is computationally heavy
Statistically low number of particles at around 20 MeV
Solution: selected only particles with +/- a few % around 20 MeV:
100 000 protons
Energy: 20±0.5 MeV
10 mrad acceptance angle
Running the simulation

## Slide 10
Simulated Bragg peak
Figure 5: Bragg peak in SmartPhantom, as simulated in BDSIM

## Slide 11
1191
∼1/84 particles enter phantom
Simulated particle depth
Figure 6: Particle depletion through simulated beamline
Quadropoles
SmartPhantom

## Slide 12
test_file.dat
Set-up
Figure 7: Side-on view of LION beamline in BDSIM

## Slide 13
GEANT4

## Slide 14
SmartPhantom in Geant4
100 mm
160 mm
140 mm
Beam
(test_file.dat)
Water
Aluminium
Kapton entry window
15 mm
Figure 8: Geant4 simulation of the SmartPhantom. Angled view (left), cross-section view (right)
15 mm

## Slide 15
Energy distribution
Figure 9: Binned energy depositions along the three axis

## Slide 16
Energy distribution
Figure 10: Binned energy depositions along the three axis

## Slide 17
k-Wave

## Slide 18
Acoustic detector array
Figure 11: 7mm diameter acoustic sensor with 300 disc elements each with 2mm diameter. Side view (left) and top view (right).

## Slide 19
proton beam
sensor array
k-Wave set-up
Figure 12: Sensor geometry and proton beam energy deposition in k-Wave simulation with 0.1mm voxels

## Slide 20
Image reconstruction

## Slide 21
Initial distribution
Figure 13: Initial pressure distribution due to proton beam energy deposition

## Slide 22
Iterative time-reversal
Model-based minimization
Back-projection
Image reconstruction methods

## Slide 23
Iterative time-reversal
Figure 14: Reconstructed pressure distribution using an iterative time reversal algorithm

## Slide 24
Model-based minimisation
Figure 15: Reconstructed pressure distribution using model-based minimisation: uses gradient descent to minimise error function consisting of the difference between the data (measured time-series) and the simulation output

## Slide 25
Back-projection
Figure 16: Reconstructed pressure distribution using back-projection

## Slide 26
SciFi Planes

## Slide 27
Scintillating fibre planes
Figure 17: Scintillating fibre planes (green) in the Geant4 simulation geometry. Off-axis view (left), side-on view (right).
4x fibre planes

## Slide 28
Scintillating fibre planes
250 μm
300 μm
50 μm
Figure 18: Scintillating fibre plane diagram (not to scale). Off-axis view of a few fibres (left) top-down view of fibres
33 polystyrene fibres
10 mm
10 mm

## Slide 29
Figure 19: Simulated energy deposited in each fibre in the four planes

## Slide 30
Analy
Bortfeld equation
Analytical approximation of dose D as function of depth in the phantom z

## Slide 31
Bortfeld fit
Use energy deposition in each plane to provide an estimate of the Bragg peak
Figure 20: Fitting Bortfeld equation to the simulated energy deposition in the 4 planes

## Slide 32
Comparison
Do the planes sample the dose distribution well enough to accurately provide a measure of the peak?
Figure 21: Bortfeld fit using the fibre planes (black), kWave reconstructed distribution (orange), “real” proton energy distribution (blue)
Energy density [MeV/cm3]

## Slide 33
Fibre plane construction
10mm
Grooves for fibres
Cavities for epoxy resin
Backing plate
Plastic frame
Figure 22: Fibre plane mid-construction
Figure 23: Fibre plane mid-construction side-view
Figure 24: Plane mid-construction with 33 fibres in place
Blu-tack to hold fibres during construction
Scintillating fibres
1 ”top frame” attached to secure fibres

## Slide 34
Finish scintillating fibre plane construction and quality assurance tests
Test the planes with a beam
Validate the simulations with experimental data
Alternative option: liquid scintillators
Topas: acoustic signals for bone, tissue, blood etc.
Future work