WP5 Consultations Update
Tony Price
LhARA 2nd Collaboration Meeting
8th Feb 2023

---
First consultation meeting
- Organised in collaboration between WP1.5 and WP3
- Originally to be in person at Charing Cross with mixed mode possible
- Train strikes meant we went fully online
- Extended invitations wider across Europe and ended up with 50+ registrations

---
Open Forums
- Afternoon not as well attended as morning sadly
- BUT those that did attend allowed very useful and informative discussions on both in-vivo and in-vitro endstations
- Ruth took nice notes during the meeting, Tony needs to write these up still! These slides are mostly a dump of Ruth’s hard work stolen by Tony
- LhARA will still have a vertical low energy and horizontal and vetical high energy endstations as per baseline designs

---
- The low energy in-vitro has an energy of 15 MeV protons. What properties are important?
  - 15 MeV +/- 2% stated
  - Question regarding other ions, possible in low energy in-vitro?
  - 35 mm diameter beam would be sufficient for studies
  - Question: are smaller beams possible on the low energy endstations?
  - Important to have sample handling as well as potential sample rotation
  - Sarah Patch (USA) asked regarding the energy spread and if this translates to the same spread at high energy stations. This will impact the acoustic signal.
  - Stress confinement discussions for the time structure of beam and dosimetry from the acoustic signals

---
- The vertical beam lines are envisaged to allow multiple hours of continuous radiation. How many samples would you irradiate in a single session? How many are too many, what is the minimum limit?
  - How many samples can you accommodate “throwing away”?
  ◦ Minimise - <10%
  - Put film on back of dishes so aware of dosimetry uncertainty/error.
  - Real-time dosimetry/monitoring allow researcher to decide to add a spare sample into the experiment. Mentioned WP4 efforts
  - Question: What is the variation across the 35mm beam? What is variation between pulses?

---
- How precise would you need the dosimetry to be during your studies?
  - 5% seems reasonable, 10% might be ok for experiments due to uncertainties in radiobiology
  - Clinical use requires overall within 5% so dosimetry 1-2% (as for photons).
  - Graphite calorimeter measurements indicates 1% is around the limit.
  - For radiobiological measurements if at least know what dose was delivered can account for that in analysis
  - Could be problematic if experiment is repeating measurements at dose level – 5% variability fine.

---
- What facilities for cell preparation, culturing, imaging, testing do you require for your studies? How close do these need to be to the irradiations? Both location and time considerations are important.
  - Do you want to irradiate one sample and look at that before irradiating whole set?
    - No answer from this one
- Need to dig deeper into temporal requirements.
  - Few comments on this one, and focusing on what the temporal properties of LhARA really offer
- Level of automation – think of level of car production line.
  - Think of getting the samples out of the room so can analyse while continuing irradiation. IRRAD facility @ CERN can remove samples robotically and continue irradiating so it is possible
  - Sample transport in and out of a shielded area can be done; just needs the requirements laying out for protection / feasibility

---
- Would you prefer the same cell dishes if irradiating on the low energy and high energy in-vitro beam lines?
  - JP - consistency; KK – able to use commercially available plates/dishes selected as appropriate for study
  - Lower energies and carbon-ions need thinner dishes to ensure Bragg peak is within the cell.
  - Probably both – some experiments want to use same, others not. At lower energies may need bespoke dishes.
  - What is needed will depend on the experiment? What biological experiments do we want to do?

---
Access comments
Is there a requirement (UK/Europe/Worldwide) for the proposed high number of beam hours?
There were questions on how users finance their beamtime, costs, etc. to see if it would be viable to experiments
Uptime service level of 95% seems too high to users when need to allow for QA etc.

---
In-vivo:
- Just rodents?
- Need for larger animals for veterinary studies? Does this restrict location for ITRF?
- In-vivo with larger animals is more translatable.
- Combination of radiotherapy and imaging so more complex end-stations
- Some contrast agents (iodine, barium sulfate) can change the depth of irradiation
- Use of window chambers on animals
- Licensing - by Home Office (no proton centre licensed yet)
- Requirement for soundproofing for animals
- Diurnal considerations (depends on waking time and experimental time on animals)
- Once irradiated, do the animals stay housed in facility? Legal requirements?
- Piggyback on back of existing animal studies
- Get RADNet/NCRI/ARR to advise, and to build a user base
- Requirement to carry out other research procedures, e.g. surgery
- Can be difficult to transport animals between facilities
Much to still consider here!!

---
Future Plans
- TP to write up notes and circulate to attendees, asking for responses to questions not covered
- Update requirements document on overleaf
- Organise next three meetings
  - 1) high energy in-vivo and in-vitro
  - 2) instrumentation
  - 3) limitations and restrictions
- Next meeting likely to be in March, location and topic TBC. We need to get organising this.